## Slide 1: Detector Response Simulations
Matt Wilson
23/10/13

## Slide 2: Introduction
Theory of Charge Induction Efficiency
Application in COMSOL
Example results from COMSOL
Combining with GEANT output
Example results of detector response
Next Tests and Information Needed
Information for reconstructions?

## Slide 3: Charge Induction Efficiency
CIE = Charge Induced/Charge Generated.
Depends on:
  Position of interaction relative to the pixel
  Trapping of charge as it drifts through the detector – E, mobility, time, detector material & defects.
  Diffusion of charge cloud as it drifts – Mobility, time.
  Pixel Geometry – Schottky-Ramo, pixel size, thickness  weighting potential
Can solve equations for these factors to build up a map of CIE of every position in the detector
  Qpix= Qdep(x,y,z,t=0) * CIE(x,y,z,tPreAmp)

## Slide 4: Charge Induction Efficiency
T.H.Prettyman, NIM A, 428 (1999) 72-80 and 232-237
Shockley Ramo:
Charge on pixel k for x=electrons (n) or holes (p)
q = charge of carrier
Number density of charge carriers
Velocity of charge carriers
Weighting Potential of pixel k at the position of the charge carriers
Changing with time

## Slide 5: Charge Induction Efficiency
For the example x=n  electrons: Number density changes with time as:
Drift of electrons in E field
Mobility of electrons in detector material
Diffusion of electron cloud
Trapping of electrons
Generation or de-trapping
Lifetime of electrons in detector material

## Slide 6: COMSOL
Need to calculate constants:
  Geometry  Easy CAD tools to build and mesh for FE.
  Electric Field  AC/DC – Electrostatics
  Weighting Potential  AC/DC - Electrostatics
Calculate charge carrier concentration at all positions t=0 to t pre-amplifier, solving:
  Drift & Diffusion = Transport of Species – Convection & Diffusion  Specify velocity and Diffusion Constant
  Trapping  Add equations
  Generation  Add equation
Combine to get CIE(t) map

## Slide 7: Example Results
Electric Potential and Weighting Potential

## Slide 8: Example Results
For 400µm thick Si – CIE quickly reaches maximum through pixel voxel in short time (<100ns).
-200V
50µm pixel pads
60µm pixel pitch
5x5 simulated
Central Pixel Solved

## Slide 9: CIE per Pixel
Use CIE solved for one pixel and re-allocate to other pixels – doesn’t account for edge effects or defects.
Chart: Position (um)
After 1µs total CIE = 1
Charge sharing between pixels accounted for.

## Slide 10: Detector Response
Load CIE(x,y,z)
COMSOL
Geant
Extract Event – x,y,z,E
W=3.62eV for Si
Is (x,y,z) in centre of pixel (CIE = 1)
No
Yes
Extract CIE(x,y,z) for Pixel N
Extract CIE(x,y,z) for Pixel N and neighbour(s)
For signal shape t=0tPreAmp
Or for total signal t=tPreAmp
Convert to Vsignal(tPreAmp) or *f(t)PreAmp
Add ASIC noise
Is V > Threshold
Yes
Information Lost
No
Store Voltage Output per pixel
Reconstruct…
Load E (x,y,z)

## Slide 11: Example Results
3.3cm
0.75mm
5cm
5cm
Marc’s geometry (?) in GEANT simulations
(0,0,0)
60µm pitch pixels.
50µm pad.
10µm gap.
833*833 pixels = 700k pixels
t = 1µs unless otherwise stated.

## Slide 12
e- track

## Slide 13: Checking it works…
Saw that Total Q pixels < Q deposited from GEANT4.
Made many changes
  Nullified noise in E  Q
  Removed trapping and generation terms
  Extended Pre-Amp time
  Made threshold of event  0
Ended up with Total Q pixels > Q deposited from GEANT4!

## Slide 14: Meshing
Chart: Position (um)

## Slide 15: Meshing
Concluded that I needed finer meshing and minimise discontinuities…
  Had to move to a Unix workstation with more computing power
  Had to buy a new license to do this
    But anyone can log onto it and use it
    Even external students if the paperwork is filled out!

## Slide 16: Meshing
Just started to test with finer mesh and no discontinuities.
Area | Min Mesh Size (µm) |  |  | Max Mesh Size (µm) |  | 
Mesh | 1 | 2 | 3 | 1 | 2 | 3
Grey | 1.1 | 1.1 | 0.15 | 26.3 | 26.3 | 15
Green & Blue | 0.15 | 0.1 | 0.1 | 15 | 10 | 5.0
Red | 0.1 | 0.1 | 0.1 | 5.0 | 5.0 | 5.0

## Slide 17: Mesh
Chart: Position (um)

## Slide 18: Meshing
Chart: 
Mesh 3 much better but still a loss of 2% in inter pixel region
Also some fluctuation at edge
Going to go to finer mesh and keep constant in inter-pixel region

## Slide 19: Summary
Have got COMSOL solving CIE but some polishing and verification to be completed.
COMSOL on work station with option for external access.

## Slide 20: Future – to discuss
Finer Mesh – Matt to complete tests
Solving output of the detector – what does Dan need?
Different geometry detectors – agree which to simulate in GAMOS and then solve CIE map.
Is solving CIE or CIE x E something that fits into Amina’s PhD?
Poster for UCL Pre-Clinical Meeting?
Approaching clinicians?
Funding options?

## Slide 21: END

## Slide 22: Energy Loss – with variation
%%%% |  |  | x | y | z | E | t
Scatter | 1 | 2 | -33.7324 | -10.7811 | 18.479 | 2.14983 | 0.130842
Scatter | 1 | 2 | -33.7327 | -10.7812 | 18.4788 | 2.60322 | 0.133246
Scatter | 1 | 2 | -33.733 | -10.7812 | 18.4788 | 3.63101 | 0.133248
%%%% |  |  |  |  |  |  | 
Scatter | 1 | 2 | -33.6823 | -0.23716 | -10.5362 | 44.5841 | 0.115776
%%%% |  |  |  |  |  |  | 
Scatter | 1 | 2 | -33.5533 | -14.343 | 1.59547 | 3.10819 | 0.12028
Scatter | 1 | 2 | -33.5717 | -14.3493 | 1.57323 | 13.4672 | 0.121843
Scatter | 1 | 2 | -33.5766 | -14.3517 | 1.5664 | 5.70247 | 0.121995
Scatter | 1 | 2 | -33.6033 | -14.371 | 1.53299 | 21.0317 | 0.122044
Scatter | 1 | 2 | -33.6052 | -14.373 | 1.53025 | 7.67459 | 0.122314
Scatter | 1 | 2 | -33.6058 | -14.3736 | 1.52926 | 1.82262 | 0.122338
Scatter | 1 | 2 | -33.6063 | -14.3739 | 1.52816 | 1.58223 | 0.122346
Scatter | 1 | 2 | -33.608 | -14.3744 | 1.52515 | 2.80993 | 0.122354
Scatter | 1 | 2 | -33.6082 | -14.3744 | 1.5248 | 1.26002 | 0.122377
Scatter | 1 | 2 | -33.6259 | -14.3762 | 1.49838 | 24.7648 | 0.122379
Scatter | 1 | 2 | -33.6313 | -14.3779 | 1.48812 | 13.7659 | 0.122588
Scatter | 1 | 2 | -33.6315 | -14.3782 | 1.48786 | 3.78164 | 0.122677
Scatter | 1 | 2 | -33.6375 | -14.3867 | 1.48268 | 29.0022 | 0.122682
Row Col # Energy(eV)
238 725 1 7.4868
413 241 2 43.706
177 441 5 34.208
178 441 5 10.193
177 442 5 58.927
177 443 5 5.2055
178 443 5 11.401
8.383keV
44.584keV
119.934keV
129.739keV
Data with variation in the charge generation and a 5keV threshold for a valid event.
Geant Output
After CIE and Simulation
Difference of Ein to Eout up to 10%

## Slide 23: Energy Loss – no variation
%%%% |  |  | x | y | z | E | t
Scatter | 1 | 2 | -33.7324 | -10.7811 | 18.479 | 2.14983 | 0.130842
Scatter | 1 | 2 | -33.7327 | -10.7812 | 18.4788 | 2.60322 | 0.133246
Scatter | 1 | 2 | -33.733 | -10.7812 | 18.4788 | 3.63101 | 0.133248
%%%% |  |  |  |  |  |  | 
Scatter | 1 | 2 | -33.6823 | -0.23716 | -10.5362 | 44.5841 | 0.115776
%%%% |  |  |  |  |  |  | 
Scatter | 1 | 2 | -33.5533 | -14.343 | 1.59547 | 3.10819 | 0.12028
Scatter | 1 | 2 | -33.5717 | -14.3493 | 1.57323 | 13.4672 | 0.121843
Scatter | 1 | 2 | -33.5766 | -14.3517 | 1.5664 | 5.70247 | 0.121995
Scatter | 1 | 2 | -33.6033 | -14.371 | 1.53299 | 21.0317 | 0.122044
Scatter | 1 | 2 | -33.6052 | -14.373 | 1.53025 | 7.67459 | 0.122314
Scatter | 1 | 2 | -33.6058 | -14.3736 | 1.52926 | 1.82262 | 0.122338
Scatter | 1 | 2 | -33.6063 | -14.3739 | 1.52816 | 1.58223 | 0.122346
Scatter | 1 | 2 | -33.608 | -14.3744 | 1.52515 | 2.80993 | 0.122354
Scatter | 1 | 2 | -33.6082 | -14.3744 | 1.5248 | 1.26002 | 0.122377
Scatter | 1 | 2 | -33.6259 | -14.3762 | 1.49838 | 24.7648 | 0.122379
Scatter | 1 | 2 | -33.6313 | -14.3779 | 1.48812 | 13.7659 | 0.122588
Scatter | 1 | 2 | -33.6315 | -14.3782 | 1.48786 | 3.78164 | 0.122677
Scatter | 1 | 2 | -33.6375 | -14.3867 | 1.48268 | 29.0022 | 0.122682
Row Col # Energy(eV)
238 725 1 7.5331
413 241 2 43.671
177 441 5 34.326
178 441 5 10.202
177 442 5 58.885
177 443 5 5.1527
178 443 5 11.445
178 444 5 5.014
8.383keV
44.584keV
125.025keV
129.739keV
No variation in charge generation…BUT there is still an “ASIC” threshold of 5keV if data is saved
Slightly closer Ein to Eout but still ~10%

## Slide 24: Energy Loss – Threshold
Too low – not saved
Saved
Example – Event 1
Charge sharing on the neighbour with one pixel having < Threshold is the major cause of the energy loss.
This effect will get worse with smaller pixels and thicker detectors  need to complete simulations to quantify.
Shows importance of an ASIC having low energy operating range.

## Slide 25: Energy Loss – Lower Threshold
Row Col # Energy(eV)
237 725 1 0.8419
238 725 1 7.5331
413 241 2 43.671
177 441 5 34.326
178 441 5 10.202
177 442 5 58.885
177 443 5 5.1527
178 443 5 11.445
178 444 5 5.014
Energy threshold  <1keV all same
0.009keV lower
0.913keV lower
125.025keV
4.749keV lower
Lowering threshold only recovers correct energy in some events  there is still a shortfall in the induced energy  is CIE correct? Is Collection time (1us) long enough in this case? Is it from trapping added to CIE?

## Slide 26: Longer Collection Time
Same with t=2us.

## Slide 27: Remove Trapping and Generation
Make lifetimes longer to negate trapping
Small
change

## Slide 28: Overlap of CIE
Chart: Position (um)
CIE overlap into next pixel  added to code - ~2% error

## Slide 29: Finer Mesh in COMSOL
Chart: Position (um)

## Slide 30: Finer Mesh in COMSOL