PAGE  
1
Draft version 1.4 – June 2004


AGATA Hybrid Preamplifiers with Pulser

Contents

1Contents


1Introduction


2Choice of the input FET


2Circuit architecture


4Active fast reset


5Programmable Pulser


7Size and pinout (segment preamplifiers)


9Size and pinout (core preamplifier)


9Target specifications


10Cables and interface specifications


15Costs




Introduction

Each Germanium crystal of AGATA will be read out by 36 segment preamplifiers + 1 core preamplifier, connected to flash ADC’s through 5-m shielded-twisted-pair cables. The role of preamplifiers is crucial to achieve a good system performance and preserve pulse-shape integrity. A list of requirements for the preamplifiers is shown in Table 1.

Table 1
	Type
	Description

	LN
	Low noise (gamma spectroscopy grade)

	WB
	Wide bandwidth (risetime of 20ns)

	SFT
	Short fall time (50us for the segments, 30us for the core)

	LP + LN
	Low power consumption of the cold FET (<20mW) maintaining LN

	HS
	High Stability of the gain

	DO
	Differential output buffer

	RDT
	Reduced deadtime, even with a background of energetic particles (10 to 100 MeV)

	CO
	Compactness


Solutions to address each of these issues have been discussed in the preamplifier group meetings. Segment preamplifiers are grouped with a three-per-board granularity. So, twelve triples will read out all 36 segments of a crystal. One core preamplifier is instead mounted on a single board. In future versions a calibration pulser will be installed on the core-preamplifier board, and used to inject a common test signal to all segments through the segment bulk capacitances. A new fast-reset technology is used in all preamplifiers to reduce the deadtime when the ADC gets saturated.

Choice of the input FET

The choice of the preamplifier input FET is one of the most critical issues of this development. In fact all FETs work cooled in the cryostat with a maximum power budget of 20 mW/FET. Nevertheless the FETs must provide the typical low-noise performance required in gamma-ray spectroscopy as a prerequisite. Note that the typical power consumption of one of such FETs, as observed in most Germanium detector setups, is three times as large. After tests performed in Milano and Cologne we have found that the n-channel FET model BF862 manufactured by Philips is adequate. In fact it provides a noise of 0.9keV fwhm on the pulser line while working at 120K with a drain voltage of 2 to 2.6V and a drain current of (8mA, i.e. a power consumption <20mW. The results of some of these tests (from D. Weisshar, May 2003) are summarized in Table 2.

Table 2
	Shaping time [us]
	Line energy [MeV]
	Line width [eV fwhm]

	
	
	IF1331 @ Vds=5.6V
	BF862 @ Vds=2.6

	
	0.122
	1.07
	1.08

	3
	1.17
	1.99
	1.97

	
	1.33
	2.15
	2.13

	
	0.122
	0.98
	0.93

	6
	1.17
	1.92
	1.88

	
	1.33
	1.97
	2.02


Circuit architecture

As shown in Fig.1, the preamplifiers have a three-stage structure, comprising a charge-sensing loop (which comprises the cold devices), a passive P/Z stage, and a differential output buffer.  An additional de-saturation circuitry is used to swiftly reset the preamplifier when an energetic event or a burst of piled up events put it into saturation. It is also possible to estimate the amplitude of the energetic event from the reset time (Time Over Threshold, TOT measurement).

[image: image1.emf]1st stage

2nd stage

Charge loop Passive P/Z Amplification

From 

detector

Cold part of 

preamplifier

Warm part of 

preamplifier

Schmitt trigger 

comparator

Discharge 

current

Output

Capacitance to 

be discharged to 

de-saturate 2nd 

stage

De De--saturation  saturation 

circuitry circuitry

/Output

1

-1

3rd stage


Fig. 1
The simplified schematic diagrams of the three stages are shown in Figs. 2, 3, and 4. 

[image: image2.emf]

          [image: image3.wmf]
Fig. 2. First stage of segment preamplifier (left) and core preamplifier (right). A 0.5-3pF capacitance (not shown) from the base of T3 to ground is used for loop stabilization.

[image: image4.wmf]
Fig. 3. Passive P/Z stage with gain amplifier.

[image: image5.wmf]
Fig. 4. Balanced differential output stage.

Active fast reset

A new fast reset technology is used to optionally reset the 2nd stage of the preamplifier when the input signals cross over a preset threshold. This is useful to reduce the preamplifier dead time brought about by energetic particles and photons that could hit the detector. The fast reset device can be excluded by pulling up a control input signal (see next).  An example of the signals seen when using the fast-reset mechanism is shown in Fig. 5.

[image: image6.emf]1000 1500 2000 2500 3000

0

1

2

3

4

5

amplitude [V]

time [us]

Signal seen at

2

nd

 stage output

1.17MeV

event

Cosmic ray

(~20MeV)

Fast reset

is active

[image: image7.emf]1000 1500 2000 2500 3000

0

1

2

3

amplitude[V]

time[us]


Fig. 5. Signals delivered by a prototype preamplifier with fast-reset device (from A. Pullia and R. Isocrate, Jan 2004). Left side: output of 2nd stage. The fast reset is visible on the 20 MeV cosmic ray. Right side: corresponding signal seen at the 1st stage output. Note that no reset occurs on it.

Control of the fast reset device is made statically or dynamically through a single wire called SHDN. One can handle the SHDN statically keeping it permanently up or down (see following points i) and ii) or dynamically (see point  iii):

i) Keep the SHDN low to make the fast reset work autonomously. Whenever a signal crosses over a preset threshold the fast-reset transient begins. The fast reset transient automatically stops as soon as the floor voltage is reached. The typical triangular shape shown in Fig. 5 (left) is thus obtained. The shape could become trapezoidal for very large signals.

ii) Keep the SHDN high to disable the fast reset. 

iii) Try to handle the SHDN dynamically to control the fast reset from the ADC module. At the beginning keep the SHDN high (fast reset disabled). Then, if an over-range condition in the ADC is detected (which means that the signal is too high) pull the SHDN low. This will force the fast reset to begin. As soon as the SHDN is kept low the fast-reset mechanism will be working autonomously, exactly as in i). So it will automatically stop when the floor voltage is reached. Of course if the SHDN is pulled up before, the fast reset transient will be over before reaching the floor. In any case the control cycle is over when the SHDN is pulled up, and the system is ready to begin a new cycle as soon as another over-range status occurs. 

Programmable Pulser

The Programmable Pulser (PP) is installed in a side of the core preamplifier board. It is a precision pulse generator designed to test the stability and linearity of the system, as well as to estimate the resolution of the segments. The output signal of PP is dc coupled to the source pin of the core j_FET through a 50 Ohm resistive voltage divider. The grounded resistor of the divider is as small as 1 to 3 Ohm and is housed in the cryostat. Thereafter the signal reaches each of the 36 detector segments following the path Csg(core)_HV-Coupling-Capacitor_Core-to-segment-bulk-capacitance. A block diagram of the PP is shown in Fig. 6.

[image: image8.png]Mode

Comparator

10dB

10dB

Programmable

/_5V

opt. /. 12V

Pulser

Attenuator

20dB

20dB

Out

13 Ohm





Fig. 6. Block diagram of the Programmable Pulser.

The PP consists of three stages, a programmable reference voltage (Progr_Vref), a chopper triggered by an external signal (Pulser In) and a programmable attenuator which comprises also an output buffer stage (Atten_Buf). The structure of the programmable reference voltage stage is shown in Fig. 7.

[image: image9.png]N c14

DAC
AD5541CR

N
L2
220uH
c12
MAX6225 o2
u1s ||'
ouT N Fh.
o R NC C17
5 5 NC 10pF

VRef




Fig. 7. Programmable reference stage.

The MAX6225 (or MAX6325) is a ultra high precision voltage reference with a very low temperature coefficient (7 ppm/ °C), very low noise (15 uVrms) and long-term stability of 40 ppm/1000 hr. The AD5541CR is a 16 bit serial input, voltage DAC. Data is written to it in a 16-bit word format, via a 3-wire industry_standard serial interface ( CS, Clock, Data_In ) 

The structure of Chopper_Pulser stage is shown in Fig. 8.

[image: image10.png]Ext. Generator €&

ADG751 BRM

ADB8058

L

74AHC1G08/SOT

5.1k




[image: image11.png]ADG719 BRM ADG751 BRM AD8058

R19 wo
11k
D T
PS
A 22 VoD
C25 Pulser Out
47nF vzt
1 o
| o o

Hry

Mode





Fig. 8. Chopper stage.

The Chopper builds around ADG751 analog switches, and is triggered by an external signal (all logic control signals are galvanically isolated). The trigger threshold & hysteresis are factory selected. The chopper mode of operation is programmable, i.e. a rectangular or an exponential-decay mode can be selected (the decay time is only factory adjustable in the range of 50us to 1ms). This stage comprises also a linear gate circuitry in order to select only one polarity, if the exponential-decay mode was selected. 

The programmable attenuator & buffer is shown in Fig. 9.

[image: image12.png]Pulser Out

o
ADG736 BRM |&

vo
ADG736 BRM A

w
ADG736 BRM A

JH

1

Ra4 R4S

u17 Hddq uts Sd uzs Qo
xS L N bt TN bt
[ 4

AT
im&




[image: image13.png]wiz
Acs

100nF
vo
ADG736 BRM A AD8058 ADG736? AD8058
U27A U30A & U268
u2e Sdd b '\J 10 - w57
4 > a7
i b - = i Dout
& & k 1 o
1 1 13 Ohm
V'} ad—am o
R53 R43 R38
R50 R51 R54 2k 1k 2k
AT~ * o677
i’m = SPolarity
»20dB
320dB
0dB

M 0dB





Fig. 9. Programmable attenuator.

The programmable attenuator consists of four stages, two of ‑10dB and two of  ‑20dB. The attenuation factor can be selected between 0 and 60 dB, with a granularity of -10 dB, respectively. The buffer can be manufactured in two versions, namely:

· either with AD8058 at +/- 6 V (standard) providing maximum output signal amplitudes  of  4V/ 50 Ohm,

· or with LM6172 at +/- 12 V (optionally) providing maximum output signal amplitudes of 8 V / 50 Ohm.

Size and pinout (segment preamplifiers)

In Fig. 10 and 11 the size and pinout of the triple segment preamplifiers is shown. The maximum thickness of the board (including components) is 7 mm. The lateral trimmers for P/Z adjustment are mounted as shown in the Figure to permit phisical access.

	[image: image14.png]K—— max50mm

upP

Side with components

DOWN

= max 70 mm

\3 Trimmers
]/ for P/Z adj.

—Fi





	[image: image15.png]GND SHDN R1 (Inh1) R2 (Inh2) R3 Inh GND

UP connector o o o o} ] o o o} o

(Side of outputs)
o o o o o o o o o

GND Out1 /Outl GND Out2 /Out2 GND Out3 /Out3
1

1
D1 FB1 GND GND D2 FB2 GND GND D3 FB3

o o o o o o o o o o
DOWN connector 0.05" Jt
(Side of detector) : ° ° ° ° ° ° ° ° ° °
+6V +12V GND GND GND GND GND GND -12vV -6V
>

01"






Fig. 10. Signals in parentheses (Inh1) and (Inh2) as well as R1, R2, R3, are shown in the pinout but are not implemented in this version.

[image: image16.png]PCB

tOp Down connector side of detector

bottom trimmer




Fig. 11. Bottom view of the board.
We have adopted polarized “Micro-Match” connectors from Tyco Electronics of the type “male on board”, shown in Fig. 12. The DOWN connector has 20 vias, the UP connector has 18 vias. Note that a hole must be drilled out on the motherboard so as to house the polarization plastic flange shown in Figs. 10 and 12.

[image: image17.png]


 

Fig. 12
The DOWN connector serves all connections to the detector (FET drain, feedback) and is also used to provide power supplies ((6V, (12V) and grounds. The UP connector is used to transmit the analog output signals and transmit/receive the digital service signals. The board plugs in over a fixed motherboard onto a 20-via "female on board" connector of the type of that shown in Fig. 13. A flexible motherboard with a similar female connector (18 vias) plugs in onto the upper part of the preamplifier.

[image: image18.png]



Fig. 13
A complete list and description of the symbols shown in Fig. 10  is reported at the end of this document. In Fig. 14 a photograph of a first prototype of a triple preamplifier is shown (courtesy of B. Cahan, Ganil).

[image: image19.jpg]°
<<
<o
<
()
<
(8]
<C
a

A120T/GEA/GANIL





Fig. 14. Photograph of a triple preamplifier prototype. Note that, in contrast to Fig. 10, in this prototype the P/Z trimmers (visible in white color) are not centered along the edge of the circuit board.

Size and pinout (core preamplifier)

In Fig. 15 the size and pinout of the core preamplifiers + pulser is shown. The maximum thickness of the board (including components) is 7mm.

	[image: image20.png]K———max50mm 3

 —

up

Side with components

DOWN

—F

e max 55 mm —— =




	[image: image21.png]Clk DIN MODE AT20 AT10 GND EN_PS GND SHDN

UP connector o o o o} ] o o o} o

(Side of outputs)
o o o o o o o o o

/CS Pulser GND AT20 AT10 out /out GND Inh
In 1

1
S GND GND GND GND GND GND GND FB D

o o o o o o o o o o
DOWN connector 0.05" Jt
(Side of detector) : ° ° ° ° ° ° ° ° ° °
GND +6V +6V +12V +12V -12Vv 12V -6V -6V GND
>

01"






Fig. 15.

A complete list and description of the symbols shown in Fig. 15 is reported at the end of this document.

Target specifications

In Table 3 the target specifications of the preamplifiers are shown. The real specifications will be available after completion of the necessary experimental tests. 

Table 3
	Property
	value
	tolerance

	Conversion gain
	110 mV / MeV (terminated)
	±10%

	Noise
	0.6 kev fwhm (0 pF)
	

	Noise slope
	12±2 eV / pF
	

	Rise time
	15 ns ± 2 ns (0 pF)
	

	Rise-time slope
	~0.3 ns / pF
	

	Decay time
	50 µs (segment)
30 µs (core)
	±5 %

	Integral non linearity
	< 0.025% (D=3.5V unterminated)
	

	Output polarity
	Differential, Z=100
	

	Fast reset speed
	~10 MeV / µs
	

	Inhibit output
	TTL
	

	Power supply
	±6V, ±12V
	

	Power consumption of input FET
	< 20 mW
	

	Power consumption (except diff. buffer)
	< 250 mW
	

	Mechanical dimension
	< 70mm x 50 mm x 7 mm (segment)
< 55mm x 50 mm x 7 mm (core)
	


It is worth adding a note about the output range and the power consumption of the differential output buffer. Using (6V as power supply for the output buffer, the output range, after signal half splitting on the termination resistor, is from 1.5 to 2.5 V (or 13.6 to 22.7 MeV) depending on  the type of used operational amplifiers. Using instead (12V as power supply for the output buffer the range gets doubled, at the expense of a larger power consumption. In Table 4 the output range and the power consumption of the output buffer are shown when using the LM6171 or rail-to-rail operational amplifiers.

Table 4
	
	LM6171/LM6172
	Rail-to-rail opamp

	
	PS=(6V
	PS=(12V
	PS=(6V
	PS=(12V

	Output range
	13.6 MeV
	27.2 MeV
	22.7 MeV
	45.4 MeV

	Power consumption
	60 mW
	120 mW
	>>60 mW
	>>120mW


It is clear that rail-to-rail operational amplifiers give better performances. However to date we have not individuated a rail-to-rail opamp with power consumption, noise and dynamic performances comparable to those of the LM6171/LM6172.

Cables and interface specifications 

The signals coming out of the preamplifiers are read out through 26-via lvds cables with MDR (Mini Delta Ribbon) male connectors on either side. The cable has a “camera-link” format with 11 individually shielded twisted pairs (Pairs 1, 2, ..., 11) and 4 spare pins connected to the shieldings (pins 1, 13, 14, and 26). The pin assignment is defined in Fig. 16 and Table 5.

[image: image22.png][e—0—>|

.

)

e O 1 12 3
an o da

5
o

B
AdAdd A
FERERUVRBRRABFAZRATR

T
o

{

;





Fig. 16. MDR 26-via male connector

Table 5. Pairs/pins association

	Preamp side
	Digitiser side
	Cable name
	Preamp side
	Digitiser side
	Cable name
	Preamp side
	Digitiser side
	Cable name
	Preamp side
	Digitiser side
	Cable name

	1
	1
	shield
	17
	10
	pair 3+
	8
	19
	pair 7-
	24
	3
	pair 10+

	14
	14
	shield
	5
	22
	pair 4-
	21
	6
	pair 7+
	12
	15
	pair 11-

	2
	25
	pair 1-
	18
	9
	pair 4+
	9
	18
	pair 8-
	25
	2
	pair 11+

	15
	12
	pair 1+
	6
	21
	pair 5-
	22
	5
	pair 8+
	13
	13
	shield

	3
	24
	pair 2-
	19
	8
	pair 5+
	10
	17
	pair 9-
	26
	26
	shield

	16
	11
	pair 2+
	7
	20
	pair 6-
	23
	4
	pair 9+
	
	
	

	4
	23
	pair 3-
	20
	7
	pair 6+
	11
	16
	pair 10-
	
	
	


Complete Signal List

	Type
	Name
	Source
	Type  / Format
	Quantity

	Analog
	Segment
	Preamp
	Analog signal  / 1 differential pair
	36

	
	Core
	Preamp
	Analog signal  / 1 differential pair
	1

	Digital
	Inh_A
	Preamp
	Digital Signal / +5.x V = High, GND = Low
	6

	
	Inh_B
	Preamp
	Digital Signal  / +5.x V = High, GND = Low
	6

	
	Inh_C
	Preamp
	Digital Signal / GND = High, -5.x V = Low
	1

	
	SHDN_A
	Digitiser
	Logic Level  / +5.x V = High, GND = Low
	6

	
	SHDN_B
	Digitiser
	Logic Level  / +5.x V = High, GND = Low
	6

	
	SHDN_C
	Digitiser
	Logic Level / GND = High, -5.x V = Low
	1

	
	AT10
	Digitiser
	Logic Level  / +5.x V = High, GND = Low
	2

	
	AT20
	Digitiser
	Logic Level  / +5.x V = High, GND = Low
	2

	
	EN_PS
	Digitiser
	Logic Level  / +5.x V = High, GND = Low
	1

	
	MODE
	Digitiser
	Logic Level  / +5.x V = High, GND = Low
	1

	
	Clk
	Digitiser
	Digital Signal  / +5.x V = High, GND = Low
	1

	
	DIN
	Digitiser
	Digital Signal  / +5.x V = High, GND = Low
	1

	
	/CS
	Digitiser
	Digital Signal  / +5.x V = High, GND = Low
	1

	
	Pulser In
	Digitiser
	Digital Signal  / +5.x V = High, GND = Low
	1

	Power
	+5.x V
	Preamp
	Supply for Logic Isolators except for Inh_C
	7

	
	-5.x V
	Preamp
	Supply for Inh_C Isolator
	1

	
	GND
	Preamp
	Cable shields & Supply for all Logic Isolators
	28


Note: all Digital Signals and Logic Levels across the interface will be isolated at the Digitizer using the Analog Devices part ADuM1100.

There are six identical connections for the segments, and one for the core. 

One cable serves 6 segment preamplifiers (or 2 triple preamplifiers: e.g. preampA tied to segments 1 2 3, and preampB tied to segments 4 5 6). So, 6 such cables are needed to read out all 36 segments of a germanium crystal. 

One more cable serves the core preamplifier, and the test pulse generator.
The signals in a Segment Cable

	Type
	Name
	Source
	Type  /  Format 
	Q.ty
	Position

	Analog
	Segment
	Preamp
	Analog signal   /  1 differential pair
	6
	Pairs 1 to 6*

	Digital
	SHDN_A
	Digitiser
	Logic Level  / +5.x V = High, GND = Low
	1
	Pair 7+

	
	SHDN_B
	Digitiser
	Logic Level / +5.x V = High, GND = Low
	1
	Pair 7-

	
	Reserved
	-
	For future expansion
	1
	Pair 9

	
	Inh_A
	Preamp
	Digital signal /  +5.x V = High, GND = Low
	1
	Pair 10**

	
	Inh_B
	Preamp
	Digital Signal  / +5.x V = High, GND = Low
	1
	Pair 11***

	Power
	+5.x V
	Preamp
	Supply for Logic Isolators
	1
	Pair 8+

	
	GND
	Preamp
	Ground
	1
	Pair 8-

	
	GND
	Preamp
	Cable shields & Supply for Logic Isolators
	4
	pins 1,13,14,26


*Pair i+ is positive swing, pair i- is negative swing, with i=1, 2,...,6

**Pair 10- swings dynamically, pair 10+ is tied to common ground, i.e. to pins 1,13,14,26

***Pair 11+ swings dynamically, pair 11- is tied to common ground, i.e. to pins 1,13,14,26

The signals in the Core Cable

	Type
	Name
	Source
	Type  /  Format 
	Q.ty
	Position

	Analog
	Core
	Preamp
	Analog signal  /  1 differential pair
	1
	Pair 1*

	Digital
	AT10
	Digitiser
	Logic Level  / +5.x V = High, GND = Low
	1
	Pair 2+

	
	AT10
	Digitiser
	Logic Level  / +5.x V = High, GND = Low
	1
	Pair 2-

	
	AT20
	Digitiser
	Logic Level  / +5.x V = High, GND = Low
	1
	Pair 3+

	
	AT20
	Digitiser
	Logic Level  / +5.x V = High, GND = Low
	1
	Pair 3-

	
	Clk
	Digitiser
	Digital Signal  / +5.x V = High, GND = Low
	1
	Pair 4+

	
	MODE
	Digitiser
	Logic Level  / +5.x V = High, GND = Low
	1
	Pair 4-

	
	DIN
	Digitiser
	Digital Signal  / +5.x V = High, GND = Low
	1
	Pair 5**

	
	/CS
	Digitiser
	Digital Signal  / +5.x V = High, GND = Low
	1
	Pair 6**

	
	SHDN_C
	Digitiser
	Logic Level  / GND= High, -5.x V = Low
	1
	Pair 7+

	
	EN_PS
	Digitiser
	Logic Level  / +5.x V = High, GND = Low
	1
	Pair 7-

	
	Pulser In
	Digitiser
	Digital Signal  / +5.x V = High, GND = Low
	1
	Pair 8**

	
	Reserved
	-
	For future expansion
	1
	Pair 9

	
	Inh_C
	Preamp
	Digital Signal  / GND = High, -5.x V = Low
	1
	Pair 10**

	Power
	+5.x V
	Preamp
	Supply for Logic Isolators
	1
	Pair 11+

	
	-5.x V
	Preamp
	Supply for SHDN_C Isolator
	1
	Pair 11-

	
	GND
	Preamp
	Cable shields & Supply for Logic Isolators
	4
	Pins 1,13,14,26


*Pair 1+ is positive swing, pair 1- is negative swing

**Pair i+ swings dynamically, pair i- is tied to common ground, i.e. to pins 1,13,14,26

Signal Descriptions.

Segment

Differential Analog signal pair. -2.5 V to +2.5 V.

Core

Differential Analog signal pair. -2.5 V to +2.5 V.

Inh_A, and Inh_B 

Active High 

(set dynamically by preamplifiers A and B)

This is a digital signal (TTL) that will be pulled up if any of the preamplifiers in triples A and B is undergoing a fast reset. It will be pulled down as soon as the fast reset is over. This can be used to inhibit acquisition of false events due to a reset transient. A source termination on the preamp side will be used. Please put no termination resistor on the receiver side to avoid half splitting of the TTL signal. The analog levels for InhibitAB are detector GND (low) and +5.x V (high) (coming with the cable). So please use such Power Supply voltages on the isolator on the cable side. Use instead the receiver 0V 5V PS levels on the receiver side.

Inh_C 

Active High (set dynamically by core preamplifier)

This is a digital signal (TTL) that will be pulled up if the core amplifier is undergoing a fast reset. It will be pulled down as soon as the fast reset is over.

This can be used to inhibit acquisition of false events caused by a reset transient. A source termination on the preamp side will be used. Please put no termination resistor on the receiver side to avoid half splitting of the TTL signal.

The Analog levels for InhibitC are -5.x V (low) (coming with the cable) and detector GND (high). So please use such Power Supply voltages on the isolator on the cable side. Use instead the receiver 0V 5V PS levels on the receiver side. The isolator in this case works also as a voltage translator.

The width of this signal can be used to estimate the amplitude of the energetic event that caused the saturation. Hence this signal should be tied to a precise counter (with the highest possible clock frequency).
SHDN_A, and SHDN_B

Active High Logic Level.

Switches off fast reset mechanism in triple preamplifiers A and B.

If this is pulled up the fast reset mechanism will be permanently switched off. If this is pulled down the fast reset mechanism will automatically work in each of the six preamplifiers when needed. 

The analog levels for SHDN_A, SHDN_B are detector GND (low) and +5.x V (high) (coming with the cable). So please use such Power Supply voltages on the isolator on the cable side. Use instead the receiver 0V 5V PS levels on the digitiser side.

SHDN_C

Active High Logic Level.

Switches off fast reset mechanism in the Core preamplifier.

If this is pulled up the fast reset mechanism will be permanently switched off. If this is pulled down the fast reset mechanism will automatically work when needed. 

The analog levels for SHDN_C are ‑5.x V (low) (coming with the cable) and detector GND (high). So please use such Power Supply voltages on the isolator on the cable side. Use instead the receiver 0V 5V PS levels on the digitiser side.

AT10

Active High Logic Level.

Attenuates pulser amplitude by 10dB.

The analog levels for AT10 are detector GND (low) and +5.x V (high) (coming with the cable). So please use such Power Supply voltages on the isolator on the cable side. Use instead the receiver 0V 5V PS levels on the digitiser side.
AT20

Active High Logic Level.

Attenuates pulser amplitude by 20dB.

The analog levels for AT10 are detector GND (low) and +5.x V (high) (coming with the cable). So please use such Power Supply voltages on the isolator on the cable side. Use instead the receiver 0V 5V PS levels on the digitiser side.
EN_PS

Active High Logic Level.

Switches on the Power Supply of built-in Pulser. If this is pulled up the built-in pulser circuitry is biased. If this is pulled down the Pulser Power Supply is set to 0V. This permits to save power when the Pulser is not used. 

This should not be handled dynamically by the Digitiser. The analog levels for EN_PS are detector GND (low) and +5.x V (high) (coming with the cable). So please use such Power Supply voltages on the isolator on the cable side. Use instead the receiver 0V 5V PS levels on the digitiser side.

MODE

Logic Level.

Selects the shape of built-in Pulser. If this is pulled up the built-in pulser provides a positive exponential decay. If this is pulled down the built-in pulser provides a square wave. 

The analog levels for EN_PS are detector GND (low) and +5.x V (high) (coming with the cable). So please use such Power Supply voltages on the isolator on the cable side. Use instead the receiver 0V 5V PS levels on the digitiser side.

Clk, DIN, /CS

Digital Signals (Clock, Data In, Chip Select) used to set the pulser fine gain.

These signals implement a simple, low-frequency, 3-wire interface through which the 16-bit input of DAC AD5541 from Analog Devices is supplied, which sets the fine gain of the pulser. The maximum clock frequency is 25MHz but it is advisable to use a substantially lower frequency (e.g. 1 MHz). The protocol of this interface can be found on the data sheet of AD5542. It is advisable to use only large numeric values (in the range from 50 to 100% of the full scale) for fine-gain corrections. Otherwise the long-term stability of the pulser could worsen significantly. Use instead AT10 and/or AT20 for coarse-gain setting. The analog levels for Clk, DIN, /CS are detector GND (low) and +5.x V (high) (coming with the cable). So please use such Power Supply voltages on the isolator on the cable side. Use instead the receiver 0V 5V PS levels on the digitiser side.

Pulser In 

Active High Digital Signal.

This is a digital signal (TTL) that will be pulled up to generate a positive transition of the pulser and down to reset it, or generate a negative transition edge (depending on setting of MODE). 

The analog levels for Pulser In are detector GND (low) and +5.x V (high) (coming with the cable). So please use such Power Supply voltages on the isolator on the cable side. Use instead the receiver 0V 5V PS levels on the receiver side.
+5.x V

Power supply provided by the preamplifiers for the Logic isolators except for the core Inhibit. The actual value will range between +5 and +6 V.

-5.x V

Power supply provided by the core preamplifier for the Logic isolator for the core Inhibit. The actual value will range between -5 and -6 V.

GND
All grounds, including the internal shieldings are to be connected at the preamplifier side and should not be connected to the digitiser’s ground.

FUTURE UPGRADE

In a future version the spare twisted pairs could be used as a serial transmission link (lvds) from the receiver to the triples/core preamplifiers (using a microcontroller to receive and distribute the signals).

This link could be used to set the P/Z fine adjustment of the triple/core preamplifiers.

IMPORTANT NOTES:

1) All detector GND's carried to the receiver along the cable must be connected at the receiver end to the cable shields found on pins 1, 13, 14 and 26 of the MDR connector.

2) All digital signals transmitted from the receiver to the detector should enter the cable through a source series termination resistor of 75 Ohms. Instead, no termination resistor should be put at the detector side to avoid half splitting of the TTL signal (with the possible exception of signal "Pulser In", which could be received through a 75 Ohm resistor and a Schmitt-Trigger comparator). 

3) All Logic Levels (see column "Format" in Complete Signal List) should be bypassed to ground with 1uF capacitors either at the transmitter or at the receiver sides.

Costs

To be filled...





























































A. Pullia - University and INFN Milano - for the AGATA preamplifier group


