Specification of the AGATA three way cryostat

General

The AGATA cryostat (see Figs. 1 and 2) will host three encapsulated, 36-fold segmented HPGe detectors of hexaconical, asymmetric shape (see Fig 3). The shapes of all three encapsulated detector types are different. The length of a detector capsule will be approx. 120mm, the diameter of the mounting lid approx. 85mm.. The detectors have to be operated at LN2 temperature. All 111 electronic signals (3 x 36 segments and 3 cores) are used for high-resolution gamma spectroscopy and position determination of the interaction points of a gamma within the detectors.

[image: image1.jpg]



Fig. 3: Schematic diagram of the three, closely packed, asymmetric, encapsulated HPGe-detectors. This unit has to be operated in a cryostat at LN2-temperature.

A shell of 60 cryostats, eachwith 3 encapsulated detectors, will be the final set up for the AGATA 4π-array (see Fig. 4).

[image: image2.jpg]



Fig. 4: A schematic diagram of half of the 4( AGATA spectrometer comprising 60 cryostats, each with 3 encapsulated detectors.

Tasks of the bidder

The bidder has to carry out following:

· Design of the cryostat for three encapsulated, asymmetric detectors in agreement with the AGATA collaboration. The mechanical constraints for the cryostat are given in more detail later in this document. The mechanical design of the cryostat has to be approved by the AGATA collaboration.

· Manufacture of the cryostat, after the design is approved by the AGATA collaboration, including obtaining all the necessary components and materials (end cap, Dewar, mechanics of internal cooling structure, feed throughs, connectors, internal cabling, HV-modules). Production of a report that shows that the mechanical specification has been met. 

· Design of the cold part of the preamplifier electronics as specified in this document.  Design of motherboards which host the warm part of the preamplifiers will be delivered from the AGATA collaboration. Preamplifiers are specified in the annex of this document.

· Assembly of the complete system once the component parts are available. Assembly includes mechanical construction of the whole cryostat, internal cabling, cold part of electronics and mounting of motherboards for the warm preamplifiers. A final cooling test with three asymmetric dummy detectors proves that the cryogenic and mechanical specifications of the cryostat are met. 

· Commissioning of the cryostat with three encapsulated HPGe detectors provided by the AGATA-collaboration. The bidder has to demonstrate that the system is running within the specifications (mechanical constraints, energy resolution, cross-talk, cryogenic system) given in this paper.

· Delivery time is 6 months after ordering. Commissioning is to be done at availability of the three encapsulated, asymmetric detectors.

Mechanics

Each encapsulated detector will have two adjacent surfaces to its neighbours (see Fig. 3). The distance between these surfaces should not exceed 0.5mm. They should not come closer than 0.2mm. 

The thickness of the cryostat end cap in the region of the flat surfaces of the encapsulated detectors should be 1mm. The distance between these surfaces of end cap and encapsulated detector should not exceed 2mm. The end cap thickness of the front surface should not exceed 1.5mm. The material of the end cap itself should be Aluminium. The alloy of Aluminium AlMg3 or an alloy of equivalent strength should be used. 

All mechanical parts and electronics (preamplifier, plugs, cables) have to fit within the angle defined by the end cap. 
Mechanics for mounting of the cryostat in a 4π-array has to be done in agreement with the AGATA infrastructure demands. It should be possible to build up a closed 4π-shell with these cryostats as shown in Fig. 4. Mechanical drawings of the cryostat are to be approved by the AGATA collaboration and the purchaser.

The crystal temperature should be 90±8K in all attitudes of the cryostat with powered electronics. The crystal temperature will be measured at a flat surface of the capsule at 2/3 of the height of the capsule.    

The cryostat should have a holding time for the LN2 of at least 18 hours in all attitudes of the cryostat with electronics powered. A fill level measurement should be available with an accuracy of 10%.

A heat resistor mechanism with a power up to 40W should be available to heat up the cold part of the cryostat. It should be placed at the cold finger. Its input leads should not be connected to the cryostats electrical ground.
Three PT100’s should be available for temperature monitoring of the crystal, the cooled electronics and the Dewar. The two signal leads of the PT100’s should not be connected to the cryostats electrical ground. The two signal leads should be connected in pairs of two in order to allow a quadripole read out of the temperature (see Fig.5).

[image: image3.jpg]



Fig. 1: Sketch of an AGATA cryostat for three asymmetric, encapsulated HPGe-detectors with LN2-cooling. Black boxes indicate warm preamplifiers, gold boxes HV-modules. The covers of the preamplifier section are omitted in this picture.

Electronics

The detectors will be operated with cooled FETs. The warm part of the preamplifier(s) developed by the AGATA collaboration will be used. The specification of this preamplifier and the connectors of the preamplifier board are defined in the paper ‘AGATA Hybrid Preamplifiers with Pulser’ written by the AGATA preamplifier team and is attached to this paper.

The cold part of the electronics (FET, feedback resistor and capacity) and the cabling to the warm part of the preamplifier has to be developed by the bidder. The FET should be operated at 130±10K. The bidder has to assure that no cross-talk higher than 0.1% is picked up in this section. 

The AGATA preamplifiers provide differential output signals and the signals should be linked via 3M-MDR26 Camera Link Style cables to an acquisition electronic. The detailed specifications of this connector are given in ‘AGATA Hybrid Preamplifiers with Pulser’.

The preamplifiers have to work with a bandwidth of 15-20MHz in the cryostat. The energy resolution of all channels has to meet the individual specification of the encapsulated de-tectors. The electronic cross talk between channels should not exceed 0.1%.  

The basis of the specification for the performance of the detectors (efficiency, resolution, cross-talk) is such that it is no worse than obtained from the acceptance test of the individual detectors done by the AGATA community. The cooled core preamplifier has to handle a test signal as specified by the AGATA community. Details are given in ‘AGATA Hybrid Preamplifiers with Pulser’.

High voltage for detector bias of up to +5000V should be generated by three modules attached to the cryostat. The AGATA collaboration proposes the miniaturised high voltage module MHx 50 103 5 2.5 from ISEG Spezialelektronik GmbH for this purpose. The data sheet of this module is attached to this document.

Plugs for power supply are D Sub 15, two rows, male, for PT100 and fill level control D Sub 15, three rows, female, and for ISEG HV control D Sub 15, three rows, male. The connector pin assignment is shown in Fig. 5.

[image: image4.jpg]Power Supply Preamps AGATA Triple:

All three detector branches supplied by one cable!
Housing plug: D Sub 15 pin, two rows, male,

mit Filter, ~ MUtron Kat. Nr. 1091914

Pin 1
Aufsicht Pin 2
Pin 3
Pin 4
Pin 5
Pin &
Pin7
Pin8
Pin 9
Pin 10
Pin 11
Pin 12
Pin 13
Pin 14
Pin 15

-6V
-6V
-12v
Ground
Ground
+12V
+6V
+6V
-6V
-12v
Ground
Ground
+12V
+12V
+6V

PT 100 Control, Capacity Control
D Sub 15 housing, three rows, female
Pt 100 in 4 pol. measurement,
capacity switchable

ISEG HV Supply and Controls
all three supplies via one plug,
D Sub 15 housing, three rows, male

Pin 1
Pin 2
Pin 3
Pin 4
Pin5
Pin 6
Pin7
Pin8
Pin 9
Pin 10
Pin 11
Pin 12
Pin 13
Pin 14
Pin 15

HVa+5Vv
HV aV Set
HV a Ground
HV aV mon
HV almon
HVb+5V
HV bV Set
HV b Ground
HV bV mon
HV bImon
HVc+5V
HV c V Set
HV ¢ Ground
HV ¢V mon
HV clmon

Pin 1
Pin 2
Pin3
in 4
Pin5
Pin 6
Pin7
Pin8
Pin 9
Pin 10
Pin 11
Pin 12
Pin 13
Pin 14
Pin 15

1. PT100 al
1. PT100 a2
1. PT100 b1
1. PT100 b2
Ground
2.PT100 al
2.PT100 a2
2. PT100 b1
2. PT100 b2
Ground
+5V Cap. Mess.
Switch
outl Capacity
ou2 Capacity
Ground




Fig. 5: Pin assignments for connectors power supply, PT100 control, fill level (capacity) control, ISEG high voltage control.





 ‘AGATA Hybrid Preamplifiers with Pulser’
Data sheet MHx 50 103 5 2.5 from ISEG Spezialelektronik GmbH 
[image: image5.jpg]AGATA Triple Cryostat - Prototype and Demonstrator
- General Specification -

Dewar LN2 Fill Port:
inner clear bore: 25.0 mm
outer thread: 1 3/8 -24 THD

Dewar with Fill Level Read Out
Outer diameter 25 cm

Content min. 4.5 litre

Holding Time ca. 18h in all attitudes

Central Body with Back Flange;
stainless steel, to be done in
arrangement with the Darresbury
and Padova Engineering groups
for the needed stability of the
complete frame structure,

Output Plugs;

3M MDR26 Camera Style plugs

for Preamps

15 Pin D-Sub for Power Supply
15 Pin D-Sub for Controlling

YT~

High Voltage Units; 3 individuel units;
_| in agreement with the AGATA Infractructure Group

Warm Preamplifiers: 3 Core preamps; 108 Segment preamps,
as defined by the AGATA Detector and Electronic Group

)~

Cold Frame Structure: Cooling of three detectors
and their cold preamps; detector operating temperature range:
90K +/- 8K for all attitudes

End Cap: defined by the shape of the used capsules;
Important: all parts of the cryostat are lying in between the
opening angle of this end cap!

i

. i :
/ ! PT 100 at 2/3 capaula halght





