Draft 9a of document produced February 2005 Last revised 14th Feb 2005


AGATA Pre-processing Hardware.

1. Contents

11.
Contents


22.
Introduction


23.
Overall design philosophy


54.
Functionality and interface specification


75.
Advanced Switching Backplanes


86.
Implementation


86.1.
Overview


86.1.1.
Functionality


96.2.
Mezzanine Implementation


106.3.
Core and segment Mezzanines


106.3.1.
Core and segment Mezzanine card input stage:


106.3.2.
Core and segment mezzanine processing:


116.3.3.
Slow control:


116.3.4.
Self Test and diagnostics (mezzanines):


116.3.5.
Data readout from mezzanine to carrier:


126.3.6.
Mezzanine to carrier interfaces


126.4.
GTS mezzanine:


126.5.
Trigger/Control Sequence:


146.6.
Diagrams of Mezzanines


146.6.1.
Diagram showing segment mezzanine


146.6.2.
Diagram showing Core mezzanine


156.6.3.
Diagram showing mezzanine optical links to/from digitiser


156.6.4.
Diagram showing GTS Mezzanine


166.7.
Carrier Implementation


166.7.1.
Introduction


166.7.2.
Crate description


176.7.3.
Card structure


176.7.4.
Acquisition Data Path


186.7.5.
Trigger system


196.7.6.
Clock management and reset system


196.7.7.
Hardware management


196.7.8.
Local FPGA Management


196.7.9.
Power supply management


196.7.10.
Slow control


196.7.11.
Power supplies


206.7.12.
Self Test & Diagnostics


206.7.13.
Embedded software


206.7.14.
Carrier diagram


217.
Summary of connections


248.
Maintenance


259.
Approximate Costings


2610.
Timescale and manpower


2710.1.
GANT Chart


28APPENDIX A     Advanced Switching for the AGATA readout chain




2. Introduction

The job of the AGATA pre-processing hardware has been discussed and refined in several meetings and this document reflects the situation following the June 2003 discussions in Munich. In essence the role of the pre-processing system is to take data from the digitiser system, extract all the useful data which can be calculated on a per-channel basis in real time, and pass on these parameters, along with the leading edge of the digitised trace from the incoming pulse, to the Pulse Shape Analysis (PSA) system. The pre-processing also interfaces with the Global Trigger and Clock system from which a clock for the digitiser and timestamp information is derived. In some cases the front end data rate might  exceed the processing capacity in the PSA, or one of the following data acquisition systems (tracking, event building or tape server). The global trigger mechanism can be used to reduce the front end data rate in these cases by pre-selection based on criteria such as multiplicity (number of active crystals) or coincidence with ancillary detectors or coincidence with beam pulses. The Global Clock and Trigger mechanism also provides the path for control and status reporting. The pre-processing hardware receives and transmits all control/status data for both itself and its associated digitiser.

The overall philosophy is explain in a little more depth in the next section. The functionality  and interface specifications are defined in section 4. Initially a phased implementation (mezzanines mounted first on CPCI carriers, then later on ATCA/AS carriers) was planned so as to make best use of available technology which will be current when  the AGATA demonstrator system is built (and later when the full AGATA system is built). This approach allowed us to defer the decision to use ATCA until we could see whether it gained market acceptance. In this context the emerging technology of Advanced Switching backplanes is a key element and is allocated a section of its own. In December 2004 the proposed 2 stage approach (CPCI, then PCI Express) was replaced and a decision made to go straight to PCI Express because delays had taken us past the planned ATCA decision date (mid 2004).

This document is the final draft. It has been revised following the meeting in Orsay (3/9/03) and revised again following discussions at the AGATA week in Italy (15-19 September) and after email discussions since September 2003. It takes into account the December 2004 decision to go straight to ATCA. The design will be frozen following the Feb 2005 AGATA week and any changes agreed during the AGATA week will be incorporated in this document.

3. Overall design philosophy

A schematic diagram is shown below. The local level processing hardware elements (digitiser, pre-processing and PSA) are shown and also their interfaces to global level processing. The pre-processing electronics is shown in green.

[image: image22.png]


[image: image23.wmf]Digitiser

36+1 outputs

Per detector

Digitiser

Digitiser

Digitiser

Digitiser

Digitiser

Digitiser

for core

Pre

-

processing

Pre

-

processing

Pre

-

processing

Pre

-

Processing 

Plus control, 

Clock, trigger

interfaces

Data

Concentrator

Pulse

Shape

Analysis

Tracking,

Event

Building,

Data 

formatting,

Online 

analysis,

Data output

Storage

(tape 

or disk)

Digitising, 

preprocessing

and PSA: repeated for each of the 120 or 180 crystals

Global trigger and

Clock system

119 or 179 

other PSA o/p


The overall design philosophy treats each crystal (core plus 36 segments) as a separate entity. Within that entity the core is treated differently to the segments. The core signal is formed by the superposition of the charge released in all the interactions in all 36 segments of the detector. So it can be used as a trigger for the whole crystal. 

The counting rate in the core contact is much higher than that in any of the segments; to a first approximation it is 36 times higher (without taking into effect the segments affected by induced charge from neighbouring segments). Since the segment electronics is triggered by the core, the rate at which the segments collect data traces is the same (high) rate as the core contact. The processing rate for traces in the segments will, therefore, be the same as in the core, although many of the traces will contain no data and will be rejected by a zero suppression algorithm. 

The need for triggering in the  core requires special interconnections with the segment electronics and also with the global triggering system. Since this connection is already made for the triggering, it will be used also as the interface point for receiving the clock, timestamps and control commands and  for reporting status. So the core contact electronics will be the master and the segments will be controlled from it. 

The fibre links from the digitiser electronics will mainly be unidirectional (transferring data from digitiser to pre-processing). However, the core contact fibre will be bi-directional (full duplex) so that clocks and control signals can be sent to the digitiser. Examples of control signals would be the offset for the baseline compensation DAC and synchronisation test pulses. One slow fibre will be provided  for each group of 6 segments to control the offset DAC at the digitiser segment inputs (and any other parameters requiring a feedback path from the pre-processing- the necessity for this link will be reviewed after the prototype tests).

The pre-processing hardware will take the incoming data streams and store traces based on the crystal level trigger information derived from the core contact (and optionally from the global trigger system too). The traces will be processed to extract parameters such as energy, time and preamplifier time over threshold (pion overload) based on the preamp “inhibit” signals sent on the top (D15) data bits from ADCs,  before passing on the leading edge of the pulse trace to the PSA, along with the parameters. Other pre-processing that could reduce PSA processing effort can also be performed if discussions with the PSA group identify any useful parameters, for example rough “seed” values might well help the iterative PSA algorithms converge to a solution more quickly and zero suppression of empty segment traces can be performed. A data concentrator is shown at the output of the pre-processing, feeding all the data from one crystal out down one data link, providing a single interface point to the PSA for each crystal.

The digitising speed will be 100MHz, so the pre-processing hardware will be designed to cope with 100MHz clock rates.

The Global Trigger and clock System (GTS) interface provides the system clock, a mechanism for broadcast/multicast commands and a trigger system. The trigger system can be used to reduce the counting rate, for example by a multiplicity filter or by a coincidence requirement with an ancillary detector or with beam pulses. Where rate reduction is not required, the pre-processing runs in triggerless mode which means that all the processed data are sent on to the PSA stage. In this case a software trigger is performed downstream, after PSA and tracking The maximum delay (latency) which can be accommodated in the pre-processing hardware while the GTS trigger decision takes place has implications for the amount of storage required in the pre-processing. It is estimated that up to 20us of trace length could be needed for processing and therefore the maximum trigger latency will be 20us. Lower values can be configured, but coincidences with more than 20us delay will be detected in software triggers.

4. Functionality and interface specification

The pre-processing hardware will contain these conceptual blocks: 

· Fibre data receivers and transceivers with serialiser/deserialiser chips to interface with the digitiser electronics

· Groups of processing elements to find, store and process traces from groups of segments (group size is implementation dependent).

· A processing element to trigger segment electronics based on observing the core contact input data stream and the global trigger interface. This part will also store and process traces from the core contact

· Global trigger/clock system interface

· Data concentrator and PSA interface

The pre-processing hardware will be reprogrammable using Xilinx FPGAs. (The optimum device family will evolve with time-  in 2003, a Virtex 2 Pro would be the best, but they will be obsolete by the time the full AGATA is built!)

In addition to the normal control interface via the Global Trigger and Clock system, there will also be an interface via the host bus system (e.g. PCI-Express) and an Ethernet control port. 
[image: image24.png]Local Level Trigger

Global Trigger

!
4
©
.
GD,
o

Digitizers
Digitizers link

Local Level Processing

PSA interface

Switching network

PSA farm




Inputs:

· 36 digitised segment signals in groups of 6 with 6 unused fibres (12 way multifibre per group of 6 segments)

· Transmitted over unidirectional multifibre connections.

· Data is converted to/from serial format before/after transmission.

· Data rate will be 80 or 100MHz, 2 bytes wide;  1.6G bits/sec of data per digitiser which requires a link data rate, including balancing bits etc. of about 2Gbits/sec. (NB this is rate is independent of counting rate- it comes from digitising 100% of the incoming signals regardless of whether they contain pulses. Thus it is independent of all physics considerations and can only be increased by upgrading the ADC.)

· 2 digitised core signals (no further spare channels)

· Transmitted over bidirectional (multi)fibre connections.

· Data is converted to/from serial format before/after transmission.

· Data rate will be 80 or 100MHz, 2 bytes wide;  1.6G bits/sec of data per digitiser which requires a link data rate, including balancing bits etc. of about 2Gbits/sec. 

· The 2 links are connected to the 2 digitiser channels set up to use different gain ranges (6/20MeV) and in the event of an error the digitiser channel attached to the working fibre will be reconfigured to the  more important gain range for the experiment in progress.
· Global Clock and Trigger interface

· Transmitted over bidirectional (multi)fibre connections.

· Data rate is <TBD>

· To be defined in discussion with GTS team (led by Marco Bellato)

Outputs

· Global Clock and Trigger interface

· Transmitted over bidirectional (multi) fibre connections.

· Data rate is <TBD>

· To be defined in discussion with GTS team (led by Marco Bellato)

· PSA Interface

· 1 data stream from the whole crystal 

· Data Rate up to 370Mbytes/sec (i.e. 3.7Gbps including balancing bits)
. If zero suppression is used then the maximum output data rate is reduced to 160Mbytes/sec
 (i.e. 1.3 Gbits/sec data over a 1.6 Gbits/sec link)

· Details of interface will be implementation dependent.

· Output Data Parameters to PSA

· Identifier word (which detector, which segment?)

· Timestamp (48 bits or less if the data is sent in timeframes)

· Event number (4 bytes)

· Leading edge of the trace (600ns)

· Energy word, and confidence/quality factor word

· Other parameters with quality/confidence factors

· Output to Digitiser

· Control commands received from Global Clock/Trigger system

· Separate fibre for pure 100MHz Clock for digitisers- no associated commands are sent on this fibre and no clock recovery is needed in the digitiser (The clock itself is recovered from the incoming GTS command serial bit stream by the GTS mezzanine)

· Offset correction values (sent via shared fibres 1 per 6 segments and via a separate fibre for the core)
5. Advanced Switching Backplanes

Conventional backplanes use a single bus over which data are transferred. The basic problem of such systems is that the data transfer bandwidth of the whole crate is defined by the backplane bus limitations. Various clever protocols have been developed to transfer data synchronously and some even send data synchronously using both clock edges. Bus widths have been increased from 32 to 64 bits and parallel local buses have been implemented (for example VSB in VME crates). These tricks allow data rates of several hundred Mbytes/sec in the fastest versions of VME and Compact PCI (CPCI). Indeed the fastest version of Compact PCI would have been sufficient for a single crystal prototype system for AGATA pre-processing. However for the AGATA demonstrator its is necessary to find something faster.

Advanced switching backplanes overcome the bandwidth problem by implementing many fast serial point to point connections, These connections set up simultaneous high speed independent data paths between processors, data sources and data sinks within a crate. So now the data transfers can all take place in parallel rather than one after the other over the single backplane. The point to point links are differential pairs and often use special high speed connectors to support the GHz serial data rates.

New concepts in Advanced switching backplanes include the topology of the point to point links and the routing protocols. The simplest topology is a star configuration with a special packet routing switch at the centre of the nodes; the most complex is a mesh where any node can talk to any other. Other topologies exist too (e.g. dual star with 2 packet switches or a ring where nodes talk only to neighbours). Packet routing protocols permit sending data packets with built-in source and destination information rather than memory mapping over a backplane. Multiple destinations and routing information can be included.

Many Advanced Switching backplanes are under development and it is important for AGATA to select the right one. After watching the market for over 1 year we have decided to go straight to the PCI Express AS switching backplane. The reason for abandoning the 2 stage proposal (initially CPCI, evolving into PCI Express, but using the same mezzanines) is that the debates about the design took longer than expected and during that period the PCI Express format has reached a sufficient level of maturity to offer a safe solution at a reasonable price without a “disposable” intermediate step.

Fuller information and a detailed  discussion of the benefits of AS backplanes for AGATA’s pre-processing are given in Appendix A.

6. Implementation
6.1. Overview

[image: image1.wmf] 


The design will be implemented using a mezzanine based approach, using CMC form factor. 

The mezzanines will be mounted on PCI Express carrier cards. They will accommodate 4 CMC mezzanines on each carrier. It is not our job to develop the ATCA infrastructure- but we have watched the market evolved as we expected to the point where the infrastructure is now (Jan 2005) available. As expected the telecoms market is adopting ATCA and so the infrastructure we need has become available as well as the security of using a standard adopted by a large commercial market. 
6.1.1. Functionality

The role of the pre-processing is to take the continuous data stream from each of the digitiser’s ADCs and extract useful information to be passed on to the PSA. 
The useful information is, as a minimum, a set of  traces corresponding to a gamma-ray’s  interaction in the detector. So, the first task is to decide which part of the incoming data stream is useful- this is achieved by running a digital trigger algorithm on the detector’s core contact data stream. When this trigger finds a pulse in the data it extracts a data trace and generates a local trigger output which indicates to all the segment electronics that they should also extract a trace from the data stream. Traces are stored locally, within each pre-processing channel. 
 At this point there are 2 options for how the pre-processing behaves. If the AGATA system’s data handling bandwidth can handle the full data from all gamma rays generated during the experiment then the pre-processing will just go ahead and put the traces in an output buffer. However, for reasons of either financial constraints in the AGATA EDAQ or over-ambitious raw gamma ray count rates it may be necessary to make a judgement on the usefulness of the pulses detected by the pre-processing, saving only the best ones. In this case the pre-processing  enquires via the GTS about whether other detectors were also active in coincidence with this one and whether the GTS system’s criteria for saving the data are met. In this mode of operation the traces are held in each channel’s local memory for up to 20us while the GTS makes a decision. Either an event reject or an event accept response is generated for each local trigger based on the GTS decision. If the EDAQ bandwidth matches the gamma ray rate then no GTS trigger decision is made and the event accept signal is generated for all detected pulses.
For events which are accepted, the pre-processing stores a trace of the digitised leading edge of the pulse from the core and all 36 segments in a buffer waiting to be sent to the PSA. 
NB- a lot of the segments will have no pulse and at this point we could save bandwidth by an intelligent zero suppression system. However,  it is planned that the PSA will be powerful enough to make its own, more intelligent, zero suppression decisions and accept the full pre-processing bandwidth.

In addition to selecting useful portions of the incoming data steam using a trigger algorithm, the pre-processing will also apply other algorithms to the data streams.
The first of these algorithms is the Moving Window Deconvolution (MWD) algorithm to determine the gamma ray energy by filtering the incoming pulse digitally. In view of the high counting rates (at least on the core), it is expected that MWD with adaptive shaping will be used- this is more complex than the fixed shaping times in existing FPGA implementations of MWD. 
The second algorithm is associated with the preamplifier’s inhibit signal. The inhibit is activated when the preamp is saturated by a pion (much higher energy than the gamma rays). The preamplifier recovers by injection of charge from a constant current source, so the length of the recovery time (i.e. the width of the inhibit pulse) is directly proportional to the energy deposited by the pion. The Inhibit from the preamps is sent from the digitiser in the previously unused to ADC data bit (D15). The width of the inhibit pulse is counted by the pre-processing (n x 10ns). The ADC baseline value immediately prior to the pion is known by the MWD algorithm and so it can be subtracted from the calculated pion energy to improve the accuracy.
Other algorithms might well be useful too. It would be possible to calculate time in the pre-processing- for example the trigger algorithm might be able to interpolate between clocks to generate an improved timing relative to the local trigger clock. It would also be feasible to calculate the timing in each segment relative to the local trigger so as to be able to recover data from piled up events where the core signal sees 2 events superimposed. Often in such cases  the sets of segments active in the 2 events will be different, and the 2 events can be separated by segment timing relative to the local trigger. 
The FPGAs in the prototype will be generously dimensioned for maximum flexibility (at a very small cost penalty) so that all options can be investigated before deciding on the devices to be used in the production version.
In addition to the trigger system described above, a CFD will be implemented in the digitiser to provide a fast (latency-free) trigger for use by ancillary detectors. The CFD output will be sent to the preprocessing over a fibre and then made available to ancillary detector trigger logic from the pre-processor.
6.2. Mezzanine Implementation

The following paragraphs describe the implementation of mezzanines in more detail. There are 3 types of mezzanine- the core, segment and GTS cards. Core and segment are very similar, differing only in detail such as the input fibre receiver footprint. The GTS mezzanine is completely different, interfacing with GTS system rather than the digitiser.

6.2.1. Mezzanine Size: 

PCI-Express ATCA cards accept 4 mezzanines. Connection between CMC and carrier card is achieved by two connectors, each 114 pins (Manufacturer Mictor, distributor: Tyco).

6.3. Core and segment Mezzanines

6.3.1. Core and segment Mezzanine card input stage: 

The input to the segment mezzanine from the digitiser will be via 12 fibre multi-fibres with 6 used for data transmitted from the digitisers. The other 6 inputs will be connected at both the digitizer and pre-processing ends, but have not yet been allocated a use. They are reserved for future use. (6 way multi-fibres are not available, so we get these spare lines for free). The fibre data receivers will produce one LVDS data stream for each input. For the core mezzanine it is planned to use the same PCB as the segment mezzanine but with a different subset of components mounted, different layout for fibre receivers/transmitters and a different program in the FPGA. The fibre link between core mezzanine and digitiser will comprise a quad transceiver (8 fibres) of which 2 will receive ADC  data, generating LVDS data streams after deserialisation and 6 will be used to transmit and receive control signals at 100 Mbits/s from and to the digitizer part, for example a 100MHz clock, a Synchronisation pulse, ADC input offset and the output of the digitiser’s CFD algorithm.

6.3.2. Core and segment mezzanine processing: 

The LVDS serial data streams will be deserialised using Rocket i/o ports in Xilinx Virtex2 Pro FPGAs. Data will be clocked into FIFO memories using the recovered clock from the serial link. At the FIFO output the clock can be either the global clock as received from the Global Trigger System (GTS) or that recovered from the incoming data. The FPGAs will not only receive but also process the data, using the MWD algorithm to calculate energy on all 38 input signals. The core signals will also be examined using a CFD algorithm to look for triggers. There was some discussion as to whether a DSP is useful at this point too- it wasn’t clear that fast floating point operations will be needed in the processing so the option of adding a DSP was noted in case it is found to be necessary during the design process, but it is not included in the costings. After some preliminary design work it was found that thermal constraints on the FPGAs force the data processing system to be split into 2 packages. A first FPGA (type XC2VP7) will deserialise the data stream via the Rocket I/O and handle the slow control via the embedded PowerPC. This FPGA will just send data to another larger FPGA (type XC2VP40) which will make all data processing and send results to the carrier board. 

The output of these FPGAs could optionally be compressed if necessary to help reduce bandwidth, for example by transmitting only differences between successive samples. The core mezzanine has less channels but more processing (trigger algorithm and higher rate MWD operation) so is dimensioned the same as the segments. The main processing device chosen (XC2VP40) uses the same footprint as 3 other FPGAs. 2 are smaller (and cheaper) FPGAs (XC2VP20 and XC2VP30) and one (XC2VP50) is larger. Reducing or increasing the FPGA size if necessary when the algorithms are implemented is simply a parts list change with no PCB rework. 
In order to set the baseline correction DACs in the digitiser channels, a baseline value must be transmitted from the pre-processing to the digitiser. Two methods are possible for transmitting this data. First, for the prototype, we will add a single fibre per group of 6 segments from pre-processing to digitiser. This fibre adds cost and complexity to the system so the need for it will be reviewed during prototype tests. An alternative is to send the baseline DAC value from each mezzanine to its carrier and then on to the centre contact TX fibre (in this case only the centre contact has a TX as well as RX fibre). The baseline information is transmitted over a special bi-directional 3 wire serial interface between carriers and mezzanines which allows for time sensitive data to be transmitted (the response time of the local Ethernet is not guaranteed on the ms timescale required for baseline correction). The 3 wire serial interface is connected between carriers by the local TCLK port alongside the trigger and timestamp information. This time-sensitive link will be retained anyway in the pre-processing to provide a fast communication path between channels.

Details of the control of the baseline offset DAC must be defined soon with the digitiser group and also the preamplifier group since the preamplifier reset will change the baseline dramatically at high counting rates, and we must remain within the digitiser ADC’s safe input voltage levels.

6.3.3. Slow control: 
On the carrier the Ethernet connection (RJ45) will be connected to a media converter and then, using Media Independent Interface (MII), connected to a software switch in the FPGA power pc on the carrier card. From the carrier card, MII links will connect the Ethernet to the other mezzanines. In this way the high current external interface for Ethernet is only done once. The software switch is a standard switch implemented by Xavier Grave within his own Linux implementation (this software is completed).  4 signal pairs per mezzanine are reserved for Ethernet connected to the power PC (PPC)  embedded in one of the on-board Viretex 2 Pro chips which will act as master. This PPC will run Linux and include a simple web server for control. 32Mbytes external RAM are included for this PPC. Some of this memory can be used for the traces to be collected in diagnostic mode when the rate doesn’t matter. 16Mbytes = 8Mwords = 8.106 * 10.10-9 = 800ms.

6.3.4. Self Test and diagnostics (mezzanines): 

A JTAG port will be provided for boundary scan data path testing and also for FPGA programming. Extensive self test and diagnostics will be built in to enhance the system reliability and fault detection/correction performance. Temperature monitoring will be provided. DACs will be connected to the FPGAs to allow visualisation of data as analogue signals to help debugging. The resulting inspection line(s) will be available on the front of the mezzanines via Lemo 00 or similar connectors. A USB interface is also planned in order to have a standalone debug system.

6.3.5. Data readout from mezzanine to carrier: 

The mezzanines will not have large derandomising memory, so the design should accommodate a maximum throughput rate of twice the expected maximum average rate. Average rate is 200bytes x 6 channels x 50kHz (60Mbytes/sec) so the maximum throughput should be 120Mbytes/sec. To accommodate this bandwidth, a 32 bit data width bus will run at 100 MHz with 4 control signals (about 400 Mbytes/s).

Several options have been discussed in order to handle this rate. The use of high rate serialised data is not possible because the connectors are not designed for GHz operation. So rates of up to 100MHz are possible. One option would be to have  48 bit wide data path from the mezzanine FPGAs (16 bits/FPGA) which would transfer 300Mbytes/sec at 50MHz. These 48 bits could be serialised/deserialised onto 9 LVDS pairs of signal lines using, for example,  DS90CR481/482 which could simplify the carrier PCB routing. 

The data will be checked in the carrier’s Virtex 2 Pro to ensure consistency of the timestamps. The use of an event number in addition to the timestamp was discussed and it was found that it could give useful information to PSA, for example during synchronisation checks. So event numbers will be transmitted to the carriers by the GTS mezzanine when the GTS system provides a trigger. In triggerless operation an event number is not meaningful so a dummy event number (either locally incrementing counter in GTS mezzanine or maybe fixed data value) will be transmitted in the event number field.

6.3.6. Mezzanine to carrier interfaces 
In each of the 114 pin connectors, 10 pins are used for power supplies and their grounds, leaving 104 pins for signals:

· Data readout path,

· Trigger path ( Timestamp, local trigger…)

· Slow Control (Ethernet)

· GTS clock

· Communication signals between mezzanines(“Message” and “Broadcast”)

For full details, refers to the “Summary of connections” chapter.

6.4. GTS mezzanine: 

The GTS interface will be handled by a special mezzanine card which will interface to the GTS system and distribute trigger signals to the whole crystal. The internal (TCLK) interface of this mezzanine is described below; the external (GTS) interface is specified elsewhere (Ref?). 

The TCLK port will be implemented at the back of the crate due to limited front panel space. It will use the user defined part of the ATCA backplane. 
The GTS mezzanine will send information to/from a local TCLK (Trigger and Clock) port . The port will distribute the following  information to the other carriers and hence to the mezzanines:

· Global clock (recovered) (transmitted via dedicated connectors)

· Local (Crystal level) trigger

· Trigger Validation

· Timestamp (sent 8 bits at a time to save pins)

· Event Number

· Communication signals between mezzanines ( “Message” and “Broadcast”)

It will also receive flow control (backpressure) signals from each carrier when, for example, the core mezzanine or  the carrier buffers reach some predefined point such as ¾ full.

6.5. Trigger/Control Sequence:
The mezzanine connected to the core contact will handle the triggering and control. When it detects a possible event it generates a local trigger request which is sent to the GTS mezzanine. In response to this request the GTS mezzanine will generate a local trigger and latch a timestamp which are distributed to all the mezzanines (via the local GTS port in the GTS mezzanine). The GTS mezzanine also sends a trigger request to the GTS through the GTS hierarchy and if the trigger condition is met, it receives a trigger validation from the GTS within 20us. The trigger validation (or rejection) is passed to the carriers and back to the other mezzanines along with a timestamp identifier (matching the timestamp of the original local trigger from GTS). When running in triggerless mode the trigger validation is automatically generated by the core contact mezzanine a fixed delay after the local trigger. 

It was agreed in June 2004 to make the trigger decisions in the GTS mezzanine rather than replicating the logic in each of the segment and core mezzanines. Therefore the timestamp will not be continuously broadcast by the GTS mezzanine. Instead it will be used as a data label and sent with the local trigger (and again with the validation or reject signal). Note that the centralised decision mechanism means that explicit rejections must be sent for events which are not validated. An event number is required by the software so will be sent with validations as well as the timestamp value.  The revised system is sketched below.

[image: image2.png]bt TS+Evimn bus for Vel

BT

it TS bus fr LT

T

Local Trigger



[image: image3.png]Il Core trigger to G TS

n Local tigger from GTS
———Soooocx Timestap (tog) from GTS
& P (tag)

TS tree system considers
Local trigger andmakes
Deciston (Validate or igncre)

Bither accept. ] ¥ sidation fram TS
- Fejection from GTS

T et (o8 kvt s from GTS
Or reject the data with this tag. ¥ sidation fram GTS
B Rejection from GTS

N X N e tte) from GTS




The core and segment mezzanines are now simpler because they don’t have to keep track of time. They can only measure time indirectly by using the local trigger as a time reference and measuring relative to the local trigger. A consequence of this is that the local trigger must be distributed to all mezzanines with low skew and a fixed relationship to the 100MHz system clock. 

The actual value of the timestamp must be captured by the GTS with low jitter too, but distribution to mezzanines of the timestamp value is not time-critical because the essential time information is carried in the local trigger signal. To avoid contention when events overlap there will be 2 different 8 bit buses from GTS to the other mezzanines for the local trigger timestamp and for the validation timestamp. 

The time between triggers is limited by the physics of charge collection in the Ge to at least 1us. So we have time to transmit a timestamp(6 bytes) + event number (4 bytes) within the shortest inter-trigger gap. 

The local triggers will be time ordered, but it cannot be 100% guaranteed (yet) that the GTS will send the validations (or rejections) in strict time order. This is, however, a design objective and will be implemented if possible in GTS.
6.6. Diagrams of Mezzanines

[image: image4.wmf] 

Figure 2 

-

 

Production

 Segment Mezzanine Pre

-

processing Board 

 

 

USB2

 

Connector

 

Stand alone

 

JTAG

 

Connector

 

Stand alone

 

 

 

FPGA

 

Vir

tex

-

IIPro

 

XC2VP7 

-

 FF896

 

396 I/O

 

 

 

CPU Trace

 

Connector

 

Stand alone

 

DUAL DAC 

Inspection

 

lines

 

LSA Trace

 

Connector

 

Stand alone

 

Optic Transmitter

 

( x 1/1 )

 

7

 

SDRAM

 

32Mbytes

 

 

LINUX

 

RAM

 

64K

 

WORD

CH 1

 

 

RAM

 

64K

 

WORD

CH 2

 

 

RAM

 

64K

 

WORD

CH 3

 

 

RAM

 

64K

 

WORD

CH 4

 

 

RAM

 

64K

 

WORD

CH 5

 

 

RAM

 

64K

 

WORD

CH 6

 

 

 

 

FPGA

 

Virtex

-

IIPro

 

XC2VP40 

–

 FF1152

 

692 I/O

 

 

40

 

Datas lines

 

40

 

Carrier

 

Access

 

Control & Trigger Bus Lines

 

 

DUAL DAC 

Inspection

 

lines

 

CON

FIG

 

MEM

 

V2Pro

 

 

120 I/O

 

27

 

40

 

40

 

34

 

PPC 

405

 

PPC 

405

 

Slow 

control

 

20

 

LSA Trace

 

Connector

 

Stand alone

 

NOT USED

 

60

 

34

 

Datas lines

 

120 I/O

 

20

 

120 I/O

 

20

 

CLOCK

 

MANAGER

 

CONFIG

 

MEM

 

Spartan

 

MICTOR

 

Connectors

 

Optic 

Receiver

 

( 12/1

2 )

 

6

 

6

 

6

 

6

 

CONFIG

 

MEM

 

V2Pro

 

 

NOT USED

 

LOADER

 

Spartan3

 

(XC3S1000)

 

 

USB2 

Interface

 

NOT USED

 


6.6.1. Diagram showing segment mezzanine

[image: image5.wmf] 

USB2

 

Connector

 

Stand alone

 

JTAG

 

Connector

 

Stand alone

 

USB2 

Interface

 

 

 

FPGA

 

Virtex

-

IIPro

 

XC2VP7 

-

 FF896

 

396 I/O

 

 

 

CPU Trace

 

Connector

 

Stand alone

 

DUAL DAC 

Inspection

 

lines

 

LSA Trace

 

Connector

 

Stand alone

 

7

 

SDRAM

 

32Mbytes

 

 

LINUX

 

RAM

 

64K

 

WORD

CH 1

 

 

RAM

 

64K

 

WORD

CH 2

 

 

RAM

 

64K

 

WORD

CH 3

 

 

RAM

 

64K

 

WORD

CH 4

 

 

RAM

 

64K

 

WORD

CH 5

 

 

RAM

 

64K

 

WORD

CH 6

 

 

 

 

FPGA

 

Virtex

-

IIPro

 

XC2VP20 to 40

 

 

–

 FF1152 692 I/O

 

 

40

 

Datas lines

 

60

 

Ca

rrier

 

Access

 

Control & Trigger Bus Lines

 

 

DUAL DAC 

Inspection

 

lines

 

120 I/O

 

27

 

40

 

??

 

34

 

OPTIONNAL

 

PPC 

405

 

120 I/O

 

PPC 

405

 

Slow control

 

20

 

LSA Trace

 

Connector

 

Stand alone

 

NOT USED

 

60

 

34

 

CLOCK

 

MANAGER

 

CONFIG

 

MEM

 

Spartan

 

MICTOR

 

Connectors

 

LOADER

 

Spartan3

 

 

 

 

(XC3S1000)

 

 

Figure 4 

-

 

Production

 Core Mezzanine Pre

-

processing Board 

 

 

Datas lines

 

120 I/O

 

20

 

20

 

Optic 

Transceiver

 

( x 4/4 )

 

40

 

2

 

1

 

2

 

3

 

CONFIG

 

MEM

 

V2Pro

 

 

NOT USED

 

CONFIG

 

MEM

 

V2Pro

 

 

NOT USED

 


6.6.2. Diagram showing Core mezzanine

[image: image6.wmf] 

Figure 6 

-

 Core Optical interface 

 

 

 

 

 

 

 

 

 

Optic 

Receiver

 

( 12/12 )

 

6

 

DATA PATH

 

2 Gbit/s per Channel

 

 

 

6

 

6

 

TRIGGER PATH

 

100 Mbit/s per Channel

 

 

 

6

 

 

 

FPGA

 

Virtex

-

IIPro

 

XC2VP7 

-

 FF896

 

396 I/O

 

 

 

Rocket inputs

 

standard inputs

 

 

 

 

 

 

 

 

 

 

 

 

 

 

 

Optic Transceiver

 

( 4/4 )

 

2

 

DATA PATH

 

2 Gbit/s per Channel

 

 

 

2

 

3

 

TOP SYNCHRO

 

PULSER SYNCHRO

 

OFFSET CONTROL

 

100 Mbit/s per line

 

 

 

1

 

SYSTEM CLOCK

 

2 Gbit/s

 

 

 

TRIGGER REQUEST

 

SPARE LINE

 

100 Mbit/s per line

 

 

 

 

 

 

 

FPGA

 

Virtex

-

IIPro

 

XC2VP7 

-

 FF896

 

396 I/O

 

 

 

Rocket inputs

 

standard inputs

 

Rocket outputs

 

standard 

outputs

 

Figure 5 

-

 Segment Optical interface 

 

 


6.6.3. Diagram showing mezzanine optical links to/from digitiser

[image: image7.png]GTS Mezzanine block diagram

-<:: Virtexl PRO rre72
ey

SystemACE  Debug Gigak

Agata LLP Meeting - Orsay 24 March 2004




6.6.4. Diagram showing GTS Mezzanine

6.7. Carrier Implementation

6.7.1. Introduction

The aim of the carrier board is to retrieve physical data from the mezzanine cards that bare the actual Local Level Processing Hardware. One carrier board will hold four mezzanines. One crystal will need one core mezzanine, one GTS mezzanine and 6 segment mezzanines. Hence, two carriers will exactly fulfil the needs of one crystal. The main functions of a carrier are the following:

· Retrieve acquisition data from the mezzanines and send it to the PSA facility through the high-speed ATCA lines of the backplane.

· Provide enough buffer storage (FIFO) to cope with random event occurrence and variable data consumption rate.

· Provide a communication system between the GTS and the other mezzanines, including those located on another carrier (through the TCLK port).

· Manage a ‘clean clock’ derived from the GTS mezzanine and distribute it to the rest of the hardware.

· Provide communication between the central slow control system and the mezzanines. This includes VHDL code and software download, calibration data and procedures etc.

· Derive and distribute the low voltage power supplies from the -48 V available from the backplane.

· Provide debug and test facilities.

6.7.2. Crate description

Several types of ATCA crates can be used. The two main selection parameters are the width (slot count) and the nature of the backplane. The width can be 23 or 19 inches, containing 16 or 14 slots. The backplane can offer either a dual-star or a full-mesh interconnection scheme. No choice has been made at this time and the card prototypes will be designed to accommodate the largest possible configuration (16 slots / full mesh) (Figure 1).

Each carrier card will provide four full-duplex serial links, each comprising two “lanes”. Each lane has a 2.5 Gbit/s (i.e. 250 MB/s) of raw data transfer bandwidth. The four links will be connected as follow:

· Two links will be routed to the two central slots of the crate (logical numbers 1 and 2). These slots will hold the Hub boards of the chassis. These two connections are provided as a standard in a dual-star backplane

· Two links will be routed to the two extreme slot-positions in the crate (i.e. logical numbers 15 and 16). These links will allow for an optional enhancement of the configuration with two local CPU boards. It should be noted however that a crate with 16 slots (23 inches) would be necessary in this case. A full-mesh backplane would increase the bandwidth between carriers and CPUs but is not mandatory since the CPUs can also communicate through the Hubs.

· All slow control links are performed through the “Base Fabric” infrastructure provided by the ATCA backplane (see slow control).

[image: image8.wmf]CPU 

(OPTION)

FOR 

PrePSA

SWITCH

GbE

or

PCIexpress

16

6

8

10

12

14

4

1,5U

8U

4,5U

23’’

3 X 2 Carriers =

1 cluster

3 X 2 Carriers =

1 cluster

11

9

7

5

3

13

2

1

15

1

2

3

4

5

6

7

8

9

10

11

12

13

15

14

16

Physical

Slots

Number

Logical

Slots

Number

CPU 

(OPTION)

FOR 

PrePSA

SWITCH

GbE

or

PCIexpress

16

6

8

10

12

14

4

1,5U

8U

4,5U

23’’

3 X 2 Carriers =

1 cluster

3 X 2 Carriers =

1 cluster

11

9

7

5

3

13

11

9

7

5

3

13

2

1

2

1

15

15

1

2

3

4

5

6

7

8

9

10

11

12

13

15

14

16

Physical

Slots

Number

Logical

Slots

Number


Figure 1 Crate description 23’’ 16 slots configuration
6.7.3. Card structure

Physical aspects:

The carrier card will comply to the ATCA standard i.e. 8 U (316 mm) in height and 280 mm in depth. Four mezzanine cards (74 x 149 = format PMC) will fit on the carrier. They will be located on the front of the card in order to provide a direct access to the fibre optics connectors to the external ADC system. Connexions between carrier and mezzanines will be provided by a set of two Mictor connectors (114 pins). Two additional connectors are provided for the power supplies on the Segment Mezzanines.

6.7.4. Acquisition Data Path

The mezzanine boards communicate with the carrier through a 32 bit / 100 MHz parallel bus. An elastic buffer is provided between the mezzanines and the FPGA of the carrier through hardwired FIFO memories. Their capacity has been preliminarily defined to 256 kB. Simulations are planned in order to reduce the capacity (maybe to 0) for cost reduction purpose (these are expensive items). It should also be noted that the acquisition bus is not bidirectional and that communication from carrier to mezzanine must be done through another path (Ethernet or message/broadcast).

[image: image9.wmf]PCI Express 2.5Gb

or 

GbE

Slow control

Data acquisition

Clock

TCLK port

DATA

Acquisition

FPGA

Carrier

XC2VP30/40

CY

2DP3120

POWER

3.3V 2.5V

1,2V 1.5V 1.8V

1GbE

MOBILE

SDRAM

128Mo

Micron

OSC LOC

CPU TRACE

LSA 

connect

.

Switch

10/100

ZL50407

FIFO

RevMII

CPU serial

GMII

PHY

32 bits

32 bits 100Mhz

RevMII

EEPROM

Config

DAFC

selectMap

JTAG 

connect

.

JTAG

JTAG

SWITCH

AS91L1006

JTAG

JTAG

SRAM

256Ko

32 bits

JTAG

CPU

HUB1

HUB2

CPU1

CPU2

OSC LOC

OSC LOC

EEPROM

Config

CPU

I2C thermal

And 

power

control

FIFO

FIFO

FIFO

C

P

U

Base

Fabric

FPGA

TCLK

TCLK PORT

PCI Express 2.5Gb

or 

GbE

Slow control

Data acquisition

Clock

TCLK port

DATA

Acquisition

FPGA

Carrier

XC2VP30/40

CY

2DP3120

POWER

3.3V 2.5V

1,2V 1.5V 1.8V

1GbE

MOBILE

SDRAM

128Mo

Micron

OSC LOC

CPU TRACE

LSA 

connect

.

Switch

10/100

ZL50407

FIFO

RevMII

CPU serial

GMII

PHY

32 bits

32 bits 100Mhz

RevMII

EEPROM

Config

DAFC

selectMap

JTAG 

connect

.

JTAG

JTAG

SWITCH

AS91L1006

JTAG

JTAG

SRAM

256Ko

32 bits

JTAG

CPU

HUB1

HUB2

CPU1

CPU2

OSC LOC

OSC LOC

EEPROM

Config

CPU

I2C thermal

And 

power

control

FIFO

FIFO

FIFO

C

P

U

Base

Fabric

FPGA

TCLK

TCLK PORT


Figure 2 : logical structure
6.7.5. Trigger system

The general trigger mechanism is the following:

· The core mezzanine detects a trigger condition based on the signal of his core-electrode and issues a Trigger-Request (TR) signal to the GTS mezzanine.

· The GTS mezzanine, upon receipt of the TR signal, issues (immediately) a Local-Trigger (LT) signal, together with a locally calculated Time Stamp (TS) to all segment and core mezzanines. Concurrently, it issues a request to the Central Trigger System (outside of the crate).

· After a certain time (to be defined), the Central Trigger System sends a Validation or Abortion signal and an event number if appropriate. The command is then broadcasted by the GTS mezzanine to segments and core. The original TS is also sent at the same time, together with the event number. It is not clear so far whether the validation chronology will respect the request chronology. (Many requests may occur before the first validation comes back and different validations can have different durations).

The carrier board will provide the hardware connections between GTS and local mezzanines as well as with mezzanines located on another carrier. Buffering and latching will be adequately provided to insure that all data are dispatched with equal logical delays. A special FPGA (SPARTAN 3) is dedicated to this function. The link between two carrier cards will be performed by a bus called TCLK-PORT. This bus will be located on the back of the card in the so-called Region 3 of the ATCA backplane. A dedicated connector will provide this function, together with a custom designed backplane section on the crate. 
The message and broadcast busses allow for quick communications between the GTS system and all the mezzanines (like general start/stop abort, synchronisation). The protocol and functions will be defined in a later stage. An additional 8 bit + control bidirectional bus has been introduced for private communication between the master and slave carriers.

6.7.6. Clock management and reset system

Each carrier board has a local, autonomous clock generator. This clock will serve as a master clock to the main FPGA and the slow-control system (Ethernet and Embedded Power PC).  The clock signal issued by the GTS card will be buffered and distributed by the carrier to all mezzanine cards (using a Motorola low jitter clock manager). However, a multiplexer will allow the selection of a locally generated clock for debugging purposes (for example if no GTS board is present/working).

6.7.7. Hardware management 

(to be documented)

6.7.8. Local FPGA Management

The code for the carrier and for the mezzanines FPGA will be loaded from the shelf management system via the backplane. The provided I2C interface may be used for this purpose. The system and protocols remain to be designed but the aim is to be able to reconfigure (reload) the whole hardware configuration by software, without any physical intervention. It should be kept in mind that a complete update will involve 2500 to 3000 programmable components.

6.7.9. Power supply management

Any ATCA crate offers as a standard an advanced management system for the power supply.

(to be documented)

6.7.10. Slow control

The slow control system is entirely based on the 10/100 Ethernet protocol. The carrier board is accessed from the outside world through the Base Fabric Interface of the backplane. This interface is provided by each of the two central switch boards (logical slots 1 & 2). A Zarlink switch chipset assures the carrier to backplane interface and provides up to 8 local Ethernet busses (in the form of reverse MII interfaces). Four are used to communicate with each mezzanine and one with the local carrier FPGA. The slow control protocol is locally managed by the embedded PPC within each FPGA.

The high level layers of the protocol (application) will be defined later (Start/Stop, Calibration, Configuration, Command/Status, Debug…)

6.7.11. Power supplies

The aim of the power supply system is to convert the -48 V available from the backplane into the various low voltages required by the mezzanine boards and locally by the carrier components. Most of the voltages are generated by local DC/DC converters. The core supply voltages of the Rocket I/O interfaces are provided by linear regulators fed from the local 3.3 V.

The following table summarizes the requirements.

	Name
	Voltage
	Usage
	Current
	Generated from
	Note

	P3V3
	3.3 V
	Carrier + all Mez.
	
	- 48 V (b. plane)
	

	P2V5_RIO_x
	2.5 V
	Carrier (rocket-IO)
	
	+ 3.3 V
	(1)

	P2V5
	2.5 V
	Carrier + all Mez.
	
	+ 3.3 V
	

	P1V8
	1.8 V
	Carrier eeproms
	
	+ 2.5 V
	

	P1V5
	1.5 V
	FPGA Core
	
	+ 2.5 V
	

	P1V2
	1.2 V
	FPGA Core (Spartan)
	
	+ 2.5 V
	

	P1V5M_y
	1.5 V
	Mez. FPGA Core
	
	+ 2.5 V
	(2)

	P1V2M_y
	1.2 V
	Mez. FPGA Core (Spartan)
	
	+ 2.5 V
	(2)


(1) One individual linear regulator for each Rocket/io (total 8)

(2) One individual supply for each mezzanine

6.7.12. Self Test & Diagnostics

(to be documented)

6.7.13. Embedded software

(to be documented)

6.7.14. Carrier diagram

The following diagram shows 2 carriers plus 8 mezzanines.

[image: image10.wmf]Power 

management

Slow

control

PSA

Or 

PrePSA

Central 

trigger

Core

Digitiser

6 Segments

Digitiser

6 Segments

Digitiser

6 Segments

Digitiser

6 Segments

Digitiser

6 Segments

Digitiser

6 Segments

Digitiser

DAC

ATCA

Master

GTS

CORE

SEG

SEG

Slave

DAC

ATCA

Slave

SEG

SEG

SEG

SEG

1 

crystal


Diagram showing 2 carriers and 8 mezzanines.

7. Summary of connections  
	MICTOR CONNECTION BETWEEN MEZZANINES AND CARRIER

	Signal
	Segment dir
	Core dir
	Notes

	

	DATA READOUT PATH

	DATA[31..0]
	OUT
	OUT
	

	DATA_CLK
	OUT
	OUT
	

	DATA_RDY
	OUT
	OUT
	

	DATA_RO
	IN
	IN
	

	DATA_ACK
	IN
	IN
	

	ETHERNET MII

	COL
	IN
	IN
	

	CRS
	IN
	IN
	

	MII_CLK
	IN/ OUT
	IN/ OUT
	

	MII_DATA
	IN/ OUT
	IN/ OUT
	

	MII_RST
	OUT
	OUT
	

	TX_CLK
	IN
	IN
	

	TX_ER
	OUT
	OUT
	

	TX_EN
	OUT
	OUT
	

	TXD[3..0]
	OUT
	OUT
	

	RXD[3..0]
	IN
	IN
	

	RX_ER
	IN
	IN
	

	RX_DV
	IN
	IN
	

	RX_EN
	OUT
	OUT
	

	RX_CLK
	IN
	IN
	

	CLOCK

	SYST_CLK<P>
	IN
	IN
	Clock from GTS

	SYST_CLK<N>
	IN
	IN
	

	TRIGGER PATH

	VALIDATION[1..0]
	IN
	IN
	

	TRIGGER REQUEST[1..0]
	NOT USED
	OUT
	

	LOCAL TRIGGER[1..0]
	IN
	IN
	

	LT_DATA[7..0]
	IN
	IN
	Timestamp : 6 bytes

	LT_SYNC[2..0]
	IN
	IN
	Timestamp strobes

	TV_DATA[7..0]
	IN
	IN
	Corresponding LT timestamp + Ev. Number : 9 bytes

	TV_SYNC[3..0]
	IN
	IN
	Corresponding LT timestamp + Ev. Number strobes

	BROADCAST

	BCAST_DATA[7..0]
	IN
	IN
	Messages from GTS to all Mezzanines

	BCAST_STROBE[1..0]
	IN
	IN
	

	MESSAGE

	MSG_DATA[7..0]
	OUT
	OUT
	Messages from Mezzanines to GTS

	MSG_STROBE[1..0]
	OUT
	OUT
	

	LLP STATUS

	BACKPRESSURE
	OUT
	OUT
	Status of the mezzanine boards

	GLOBAL_RESET_1
	IN
	IN
	

	LLP_STATUS[7..2]
	OUT
	OUT
	

	GTS STATUS

	GLOBAL_RESET_0
	IN
	IN
	Status of the GTS board

	GTS_STATUS[7..1]
	IN
	IN
	

	DATA DESKEW

	DATA DESKEW[5..0]
	NOT DEFINED
	NOT DEFINED
	NOT DEFINED

	FPGA PROGRAMMATION

	TCK
	IN
	IN
	

	TMS
	IN
	IN
	

	TDI
	IN
	IN
	

	TDO
	OUT
	OUT
	

	FPGA_PROG
	OUT
	OUT
	

	FPGA_INIT
	IN
	IN
	


	FRONT PANEL CONNECTION

	SIGNAL
	DIRECTION
	NOTES

	

	SEGMENT MEZZANINE

	SERIAL DATA IN ( x6 channel )
	IN
	Optical fiber : 2 Gbits/s per channel on RocketI/O

	TRIGGER PATH ( x6 channel )
	IN
	Optical fiber : 100 Mbits/s per channel on standard inputs

	OFFSET CONTROL
	OUT
	Optical fiber : 100 Mbits/s

	ANALOGUE INSPECTION (x4)
	OUT
	From DACs to front panel connectors

	DIGITAL INSPECTION (x4)
	OUT
	From FPGAs to front panel connectors

	USB2 PATH
	IN
	For standalone debug

	

	CORE MEZZANINE

	SERIAL DATA IN ( x2 channel )
	IN
	Optical fiber : 2 Gbits/s per channel on RocketI/O

	SYSTEM CLOCK
	OUT
	Optical fiber : 100 Mbits/s

	TRIGGER REQUEST
	IN
	Optical fiber : 100 Mbits/s per channel on standard input

	SPARE
	IN
	Optical fiber : 100 Mbits/s per channel on standard inputs

	TOP SYNCHRO
	OUT
	Optical fiber : 100 Mbits/s

	PULSER SYNCHRO
	OUT
	Optical fiber : 100 Mbits/s

	OFFSET CONTROL
	OUT
	Optical fiber : 100 Mbits/s

	ANALOGUE INSPECTION (x4)
	OUT
	From DACs to front panel connectors

	DIGITAL INSPECTION (x4)
	OUT
	From FPGAs to front panel connectors

	USB2 PATH
	IN
	For standalone debug


	TCLK PORT with notes  (all signals are differential)

	SIGNAL
	GTS/CORE CARRIER
	SEGMENT CARRIER
	NOTE

	

	CLOCK

	CLOCK
	OUT
	IN
	Clock from GTS

	

	TRIGGER PATH

	VALIDATION[1..0]
	OUT
	IN
	Validation from GTS

	LOCAL TRIGGER[1..0]
	OUT
	IN
	Local trigger from GTS

	LT_DATA[7..0]
	OUT
	IN
	Timestamp : 6 bytes

	LT_SYNC[2..0]
	OUT
	IN
	Timestamp strobes

	TV_DATA[7..0]
	OUT
	IN
	Corresponding LT timestamp + Ev. Number : 9 bytes

	TV_SYNC[3..0]
	OUT
	IN
	Corresponding LT timestamp + Ev. Number strobes

	

	BROADCAST

	BCAST_DATA[7..0]
	OUT
	IN
	Messages from GTS to all Mezzanines

	BCAST_STROBE[3..0]
	OUT
	IN
	

	

	MESSAGE

	MSG_DATA[7..0]
	IN
	OUT
	Messages from Mezzanines to GTS

	MSG_STROBE[3..0]
	IN
	OUT
	

	

	LLP STATUS

	BACKPRESSURE
	IN
	OUT
	OR of all segment mezzanines

	GLOBAL_RESET_1
	IN
	OUT
	Reset from LLP

	LLP_STATUS[7..2]
	IN
	OUT
	Status of LLP

	

	GTS STATUS

	GLOBAL_RESET_0
	OUT
	IN
	Reset from GTS

	GTS_STATUS[7..1]
	OUT
	IN
	Status of GTS


TCLK PORT CONNECTIONS 

	Signal name
	Signals
	Source
	Destination
	Level
	Notes

	Trigger validation

Tv<1:0>
	2
	GTS Mez.
	Slave carrier, all Mez.
	LVTTL 2.5 (local)

LVDS (Remote)
	(1)

	Trigger validation Synchronisation

Tv_sync<3:0> 
	4
	GTS Mez.
	Slave carrier, all Mez.
	LVTTL 2.5 (local)

LVDS (Remote)
	(1)



	Trigger validation Data

Tv_data<7:0>

timestamp +

Evt-Number
	8
	GTS Mez.
	Slave carrier, all Mez.
	LVTTL 2.5 (local)

LVDS (Remote)
	(1)



	Local Trigger

Lt<1:0>
	2
	GTS Mez.
	Slave carrier, all Mez.
	LVTTL 2.5 (local)

LVDS (Remote)
	(1)



	Local Trigger synchronisation

Lt_sync<2:0>
	3
	GTS Mez.
	Slave carrier, all Mez.
	LVTTL 2.5 (local)

LVDS (Remote)
	(1)



	Local Trigger timestamp

Lt_data<7:0>
	8
	GTS Mez.
	Slave carrier, all Mez.
	LVTTL 2.5 (local)

LVDS (Remote)
	(1)



	Backpressure
	1
	Slave Carrier
	GTS Mez. 


	LVTTL 2.5 (local)

LVDS (Remote)
	(2)

	Reset from LLP

Global_Reset_1
	1
	Master

Carrier
	Slave carrier,

GTS Mez.
	LVTTL 2.5 (local)

LVDS (Remote)
	

	Status of LLP

LLP_Status <7:2>
	6
	Slave Carrier, all Mez.
	GTS Mez.
	LVTTL 2.5 (local)

LVDS (Remote)
	

	Reset from GTS Global_Reset_0
	1
	GTS Mez.
	Slave Carrier, all Mez. 
	LVTTL 2.5 (local)

LVDS (Remote)
	

	Status of GTS

GTS_Status <7:1>
	7
	GTS Mez.
	Slave Carrier, all Mez. 
	LVTTL 2.5 (local)

LVDS (Remote)
	

	Message synchronisation

Message_sync<3:0>

(sync answer to GTS. Driven by GTS)
	4
	GTS Mez.
	Slave Carrier, all Mez.
	LVTTL2.5 (local)

LVDS (Remote)
	(1)



	Message data

Message_data<7:0>

(answer to GTS)
	8
	Slave Carrier, all Mez.
	GTS Mez.
	LVTTL2.5 (local)

LVDS (Remote)
	(1)



	Broadcast

Synchronisation

B_cast_sync<3:0>

(message form GTS)
	4
	GTS Mez.
	Slave Carrier, all Mez.
	LVTTL2.5 (local)

LVDS (Remote)
	(1)



	Broadcast data

B_cast_data<7:0>

(message form GTS)
	8
	GTS Mez.
	Slave Carrier, all Mez. 
	LVTTL2.5 (local)

LVDS (Remote)
	(1)



	Message Carrier master to/from slave

msg_car<9:0>
	10 pairs
	Master Carrier
	Slave Carrier 
	LVDS
	(3)

	Clock in/out
	2 pairs
	GTS Mez. via fan-out on carrier
	Slave Carriers, all mezzanine
	LVPECL 2.5 (local)

LVDS (Remote)
	(4)


(1) Also has internal connection on master carrier to core and segment mezzanine.

(2) Logical OR of all segments mezzanine backpressure signals.
(3) Allow direct message between Master  & Slave Carriers

(4) Internal connection to mezzanine within Carrier. Transmitted to slave Carrier via backplane.
Host Interface

In addition to the connections described above which will exist for all host systems, the carrier must also have an interface to the host bus (CPCI Express(AS)). The host interface will either be the entry point to the PSA directly or via another PSA interface.

8. Maintenance 

This is a key issue and several ideas were raised:

It is important to involve one or more engineers from the host lab in the design/testing phase so that local support people feel some ownership and also understand the system.

Self test and diagnostics must be built in everywhere

We need to decide what is our repair policy- will we employ someone to do repairs? Will we employ a company to do repairs (and possibly commission cards too)?  Note that the volume of repairs in a 6000 channel system could create a significant workload.

· Deliverables must include not only hardware, but full documentation (circuit descriptions, user manuals and commented code with documentation for VHDL).

9. Approximate Costings 

	[image: image25.png]Schematic of the Digital Electronics and Data Acquisition System

for AGATA
A

36+1
Outputs

Pulse Shape

Pre-Processing Analvsi
nalysis

All other
detectors

Data

Digitiser Digitiser Readout

for core

Pre-processing

Tracki
plus control, Event

Event

Clock, trigger Building
interfaces Global trigger and || <= | Data format
clock system Online

Analysis
Data Output






10. Timescale and manpower

All the following figures assume 100% effort devoted to AGATA- we are unlikely to get this (typically 50% to 75% is realistic) so elapsed time is likely to be longer than the man months figure.

Core/Segment Mezzanine 
hardware 
9 man months 





FPGA

2x9 man months 

So with 3 people working in parallel this could be ready to test in 1 year.

ATCA Carrier 

hardware 
12 man months 





FPGA

9 man months 

Linux Software for PPC 12 man months

So with 3 (or 4)  people working in parallel this could be ready to test in 1 year although the full software will take longer with only 1 person available 50% of the time.

GTS Mezzanine
hardware 
9 man months 




FPGA

6 man months

So with 2 people working in parallel this could be ready to test in 1 year.

Digitiser Mezzanine: CSNSM Orsay (hardware and FPGA)

Carrier: IPN Orsay (Hardware, Software, FPGA) 
GTS Mezzanine: Padova/LNL (hardware and FPGA)
Timescales

Some preparatory work has been done during 2004 including discussions about interfaces with the other engineering teams and discussions about costs with AMB. The decision to freeze the design has been hard to make as a result of these discussions, which means that we have no clear start date. Assuming that design work starts now (Feb 2005) and there are no further changes to the specification. Prototype delivery dates (for the start of testing) are expected to be:

Core and Segment mezzanines (hardware and FPGA firmware)
Summer 2005 

GTS Mezzanine (hardware)





February 2005 

GTS Mezzanine (FPGA firmware)




Summer 2005 

ATCA Carrier (hardware and FPGA firmware)


January 2006

Testing of each part individually will be complete by end March 2006

Testing of all parts together (including digitiser and PSA)  will take April 2006- Summer 2006.

Manufacture of Production units (mezzanines and carrier) will start in autumn 2006 with deliveries of tested units late 2006/early 2007.

10.1. GANT Chart


[image: image11.png]2005 2006 2007
0 {Tock Neme Gz [ous[ova [owiJowz[aus[aud [oui]ouz [ous [aud |avi]avz]aus
2 |Design frozen @02/02105
3
4| Core Mezz design 02/02/05 29/03/05
5 | Core Mezz CAD input 30/03/08 ["},26/04/05
6 | Core Mezz PCB layout 27/04)0¢ 23/06/05
7 | Core Mezz assembly 44106105 [ |,21/07/05
8 | Core Mezz test 22007105 010910
g
10 | Segment Mezz design 02/02/05 29/03/05
11| Segment Mezz CAD input 30/03/08 ["},26/04105
12| Segment Mezz PCB layout 2710405 23106105
13 | Segment Mezz assembly 44106105 [ |,21/07/05

egment Mezz test 22107105 1-20/09.0
5
16 |GTS Mezz design pro4 1211004
17 | GTS Mezz CAD input 13110104 [ ],00/11/0:
18 |GTS Mezz PCB layout 10111/04 20101105
19| GTS Mezz assembly 21101/05 [ |17/02/05
20 |GTS Mezz test 18102/05 [[|-472:0305 — |
71

[ 22 |ATCA Carrier design 02002105 tﬂms
23 |ATCA Carrier CAD input 08/07/05 06/09/05
24 | ATCA Carrier PCB layout 07/09/05 15111105
25 |ATCA Carrier assembly 16111105 [ 1812105
26 | ATCA Carrier test 14112105, 23103106
77
28 |ATCA Carrier+Mezz test 131011 06104106
]
30| Test Digitiser+preproc+PSA 07/04/06 7107106
3

[ 32 |Production Mezz designs (3) 23105106 7107106

[ 33 |Production Mezz CAD input (3) 18107106 [ p4/07106
34| Production Mezz PCB layout (3) 25/07/06
35 | PreProduction Mezz assembly (3)

36 | PreProduction Mezz test (3)

37| Full Production Mezz assembly (3) 09/02/07
38 | Full Production Mezz test (3) 23/03/07
g

[

41| Production ATCA Carrier design 20/06/06

42 |Production ATCA Carrier CAD input 18/07/06

43| Production ATCA Carrier PCB layout 25/07/06 05/10{06

44| Pre-Production ATCA Carrier assembly 06/10/06 1106

45| Pre-Production ATCA Carrier test 03/11/06 /1106

46| Full Production ATCA Carrier assembly 01112106, 09/02/07
47| Full Production ATCA Carrier test 1510 2310307
[

49 stem Integration 20001707 ],20/04/07
EY l

5

Experiments start

*

Project: Timelinev1
Date: Tue 08102105

Task
Critical Task

Progress

Milestone

Summary

Rolled Up Task

Rolled Up Critical Task

Rolled Up Milestone

[

Rolled Up Progress
Split
Extemal Tasks

Project Summary

External Milestone

Deadiine

Page 1





APPENDIX A     Advanced Switching for the AGATA readout chain

Proposal of this document

· Highlight the Advanced Switching (AS) features of PCI Express technology in the context of the AGATA readout system and Pre-Processing Farm Interface.

· Propose an event reassembly protocol  that exploits AS benefits.

Relationship To Other Research & Practice 

This document is part of the Hardware Level Processing Proposal for the AGATA Experiment to be presented at the first AGATA week in Legnaro, 15-19 Sep. 2003.

A basic knowledge of the PCI Express technology is assumed.

Foreword

Without delving into deep detail of the AGATA structure, a DAQ column can be sketched as shown in the following figure:


[image: image12]
Fig. 1 – The AGATA readout column

This sort of digital pipe highlights the data flow starting from the digitizer of a single channel down to the Pulse Shape Analysis farm where position estimation is carried out. Interaction with two types of trigger is also underlined. By associating this data flow scheme to each single detector channel of which AGATA (or the demonstrator) is comprised, the modularity and scalability of the system is exploited in a two dimensional space: on the x dimension we have the detectors while on the y dimension we have the different functions to be performed on the data. By factorizing these functions into subsystems whose performance is independent of each other we can exploit the upgradeability of the system to its maximum extent. By replicating these columns we exploit scalability, according to the evolving needs (prototypes -> demonstrator -> final system). The complete columnar model of detector channel is reported in figure 2, where the interaction with the Global Trigger and Synchronization (GTS) Network is detailed. 

[image: image13.png]GTS
Time, Trigger and Control

(—o—)

Detector channel

Frontend model

TIME, TRIGGER AND CONTROL

DISTRIBUTION SYSTEM

CLOCK 100 MHz

CCRESET >
L1 TRIGGER AC

EC RESET >

1

COUNTER

FrontEnd System

DIGITAL PIPELINE

READOUT WINDOW

T

Event
couner

TEST SIGNALS—>
INITIALIZATION——

FAST
CONTROLS

—

PARAMETERS ——

(PROCESSING ELEMENT)

DERANDOMIZER

LOCAL TRIGGER PROCESSING
LOCAL PRE-PROCESSING

MULTIPLEXER (SPARSE READOUT)

PROCESSING ELEMENT
¢ FE IDENTIFIER

TIME TAG

EVENTTAG
FORMATTING

PSA LINK INTERFACE

PSA Farm





Fig. 2 – Model of a detector channel

Thus, by aggregation of m channels, the model for a hardware pre-processing unit turns out to be as in the following figure


[image: image14]
Fig. 3 – Model of a Hardware Level Processing (HLP) unit

and hence the model for a individual detector would be as the following:

[image: image15.png]sector channel

local or global

trigger
Hardware Level
Pre-Processing = =
{HLP) Card

PSAinterface

M PSAFarm





Fig. 4 – The Detector model

The target of this discussion is to illustrate how the mapping of the bottom part of this block diagram might be accomplished in a seamless way by referring to an emerging technology known as PCI Express Advanced Switching

PCI Express Advanced Switching

PCI Express is designed to provide a standards-based, high performance, highly pin-efficient scalable architecture. PCI Express increases performance and reliability required in distributed systems. When combined with the new ATCA specification, PCI Express provides not only a standardized form factor, but with the power and flexibility to meet the system requirements of computing and networking products including servers, communication products, storage systems, and industrial control. PCI Express architecture is ideal for systems that require a high speed, pin-efficient, and scalable bus. It is also targeted to address a wide array of systems, from desktop PCs to chassis based telecom systems such as storage, routers, and high-end compute platforms.

Layer 1: Rather than a shared-bus architecture, PCI Express specifies a point-to-point system. The basic data transmission unit is two pairs of LVDS wire, called a lane. The two LVDS pairs are simplex communication in opposite directions, which eliminates data collisions. PCI Express utilizes 8 bit/10 bit encoding for clocking, which eliminates issues related to clock skew as well as reduces pin count. Currently, PCI Express specifies 2.5 Gbits/sec per lane. This performance is scalable; bandwidth can be linearly multiplied by adding more lanes. The PCI Express specifications allow a number of lane widths (1,2,4,8,12,16,32) to allow not only efficient simplified routing, but also the ability to grow as bandwidth requirements increase. In addition, PCI Express is architected not only to allow for higher speeds utilizing copper, but also to use higher-speed media such as fiber for future generation systems.

Layer 2: PCI Express’ Data Link layer, provides reliable data transmission, both in terms of data sequencing, as well as data integrity. This is accomplished via sequence numbering and cyclic-redundancy check (CRC), respectively. PCI Express employs a credit-based flow control to reduce bus bandwidth waste. Credit-based flow control eliminates bandwidth-robbing polling by allowing devices to transmit data only when the receiving device has the resources to accept all of the data.

The PCI SIG, PCI Express’ governing body, developed two related standards, Base and Advanced Switching:

· Base’s primary goal is to remain driver- and OS-code compatible with PCI while providing a high-speed, low-pin-count serial connection.

· AS relaxes the code capability requirement in order to standardize advanced features required for distributed systems.

	ADVANCED SWITCHING FEATURES

	Compatible with PCI Express L1/L2
Uses same physical interface (L1) 
2.5-Gbit/s link, up to 32 links/connection
Uses same data link layer (L2)

	Different protocol layer than PCI Express Base
AS uses peer-to-peer architecture
PCI Express AS requires new applications
PCI Express Base supported through PI tunnels

	Source Routing
Switches don't require unicast routing tables
Peer-to-peer doesn't require central route management
Switches can notify source of delivery failure

	Protocol Interfaces (PI)
Allow high-performance tunnelling of any protocol, including PCI Express Base (PI8)
Used for AS features like multicast and congestion management

	Virtual Channels
Support in-order and bypass queuing model
Implement class-of-service (CoS) support

	Events
Handle congestion and power management
In-band high-availability support like hot-swap


Table 1 – Features of the AS layer

One of the more significant changes in AS is the need to work in the new high-end topologies. In order to address the need for reliable, high-speed bandwidth and processing, systems are looking less like the host-based topologies found in PCs, and more like truly distributed computing systems. CPUs in these systems cannot afford the more cumbersome host intervention. Instead, routing is accomplished not by memory mapping, but rather by path routing as shown in figure 5. 

[image: image16.png]Memory or 10 Address Space

Root complex

Li mit

—

Downstream

Example

Upstream
Inverse decode

Address Routing

Li mit

Base




Fig. 5 – Address routing in PCI Express 

Path routing, as done in the AS layer, circumvents the host control by specifying the destination relative to the switch (figure 6). This more efficient scheme allows for more efficient bus utilization and reduces latency.

[image: image17.png]4 turns

Packet #

Ingress » Packet
Switch #2 Egress
Interpret assigned
Turn Pool bits
Interpret assigned
Turn Pool bits

2 turns




Fig. 6 – AS Unicast routing avoids need of switching tables in the switches 

The AS layer natively supports broadcast and multicast communications, a key feature in the multiprocessing environment (figure 7). The AS header specifies a multicast group ID that indexes a multicast routing table implemented in the switch and managed by software.  The table specifies which ports are associated to each ID, as shown in fig.7 for multicast ID 2.

[image: image18.png]Packet

Ingress »

f
-t
L

Interpret assigned
multicast ID

(Compare in Multicast Table)

MGID

1

OUTPUT PORT

2 3

4

5

000

0

0

001

1

1
1

1

010

olofole

01
00
01

0

1

6
0
0
1

s lo|o N





Fig. 7 – AS Multicast routing 

	PROTOCOL INTERFACES

	Standard PIs
Spanning tree generation/Multicast
Congestion management (CM)
Segmentation and reassembly (SAR)
Device management
Event reporting

	Companion PIs
PI8 - PCI Express Base
Simple load/store (SLS)
Simple queue packets (SQ)
Secure data transport (SDT)

	Vendor-specific
30 reserved PIs
255 PIs total


Table 2 – Protocol Interfaces of the AS layer

Another significant addition in AS is Protocol Encapsulation. This feature allows AS Systems to tunnel other protocols intact within the AS fabric and then quickly reconstitute the protocol. This provides a high-performance, switching fabric that allows legacy systems to extend their lifespan. AS specifies other capabilities such as multicast, peer to peer, and multi-host necessary to distributed systems. 

Proposal

Going back to the model of the individual detector readout (reported here for ease of reference)

[image: image19.png]sector channel

local or global

trigger
Hardware Level
Pre-Processing = =
{HLP) Card

PSAinterface

M PSAFarm





Fig. 8 – Model of the Detector 

It’s easy to realize that this scheme is essentially a classic multiple producers – multiple consumers configuration in which the token produced and consumed is the detector event, whose fragments are distributed among some of – possibly all – the producers (the Hardware Level Processing cards). Therefore the very first task of consumers (PSA Farm) is having the single fragments of each event reassembled in their own memory queues before pulse shape processing. This is the well known event building stage.

It is straightforward to recognize that a barrel shifting approach, as shown in the following figure 9 for a 4x4 case, would map directly from theory to practice by exploiting the AS features of PCI Express.

By assigning a virtual channel (VC) from each source to every destination we establish the mesh of required connections among HLP cards and PSA farm processing units. By using the event number – a well known parameter to every HLP card, distributed synchronously by the Global Level Trigger and Synchronization Network – modulus the number of processing units in the PSA farm, we fix the upper limit of the round-robin mechanism.

[image: image20.png]Carriers barrel shifter

Each aarrier transports
event fragments to the
same BU destination

wiypuob|y ysd ol

spied dTH wou4




Fig. 9 – Barrel Shifting based on AS 

In this way the path based routing of the AS layer, as shown in figure 6 is automatically obtained from the event number without the intervention of any external agent which knows where to route the appropriate fragments. It is worth noting that the credit based flow control of the AS layer implicitly constitutes a back pressure mechanism – a vital feature of a distributed readout. In fact this type of flow control allows the delivery of packets from source to destination if sufficient memory space is assured at the recipient. This means that a congestion in one of the processing units of the PSA farm would automatically translate to a stop signal on all the HLP cards.

The readout proposed above is a push-type mechanism with an implicit feedback; actually the amount of data flow is automatically regulated by the aggregated speed of data processing in the PSA farm. 

Instead a pull-type mechanism exploits data requests in the opposite direction, from consumers to producers, but normally relies on an external control agent that keeps track of free processing units and routes events to them. 

With AS, a fully distributed approach becomes possible and no external agent is required. AS allows a peer-to-peer approach by keeping distinct the memory spaces of each single processing unit. By means of the multicast mechanism, as shown in figure 7, each processing unit of the PSA farm can notify all the sources of the will to book the readout of next event in the HLP cards queues.

In this way, all the HLP cards can keep updated and synchronized a queue of free processing units where fragments of events have to be downloaded. A transparent memory access from each of the processing units to all the HLP card queues would complete the event building. This is schematically shown in the next figure.

[image: image21.png]



Fig. 10 – Distributed Event Building based on AS 

This policy corresponds to a fair queuing rule at every machine, where requests arriving earlier are served before later arrivals. While this scheme is basically a distributed fist-come-first-serve scheduling policy, it is worth noting that any other known distributed scheduling mechanism can be implemented by relying on the native support for distributed multiprocessing of PCI Express AS hardware protocol. The most suitable solutions in the AGATA readout would be a subject of investigation, based essentially on parameters like the actual topologies used, typical loads of the network, hardware and software latencies, round trip times, constraints on costs, etc.

Conclusions

A new hardware based protocol supporting advanced switching techniques is being proposed as an extension of the PCI Express technology. Coped with a high bandwidth medium, this new protocol has potential benefits for highly distributed and interconnected systems, as in the case of the readout system of AGATA. While the concept of event building by means of switching architectures is an established practice in the nuclear and high energy physics experiments, the interest for this new protocol is manifold, due to its strong hardware support for distributed processing, the transparent distributed memory management, the scalability of bandwidth, the standardization. In the past, due to limitation in the switching structures, the most common practice was to buffer fragments at the level of sources (HLP cards) and to send them at destinations where a sort of batch event building could take place. What AS is promising instead is the possibility to make event building on an event basis and transparently through distributed memory access. This might turn out to be a great simplification on the quite complex data flow management and processing. 

The early adoption of AS is unknown, as is for the time to market of the first silicon products. Many things remain unclear, for the time being, as per the possible use of AS in AGATA and to which extent. It is not clear also to which extent the features of AS can be complemented (or replaced) by custom built features in the Base PCI Express protocol, a possibility that might prove useful due to the static nature of AGATA readout interconnections.  These reasons induce the need of an extensive investigation on the subject. 

Comments

AS offers major advantages in throughput and functionality only available in a fabric-oriented architecture. At the time of writing, the AS specification are not yet available, but are expected to be released in the third quarter of 2003 and hardware isn't expected until later in the year. The PCI Express Base standard has been released since April 2002 and the first chips are becoming available now. PCISIG (www.pcisig.com) handles PCI Express Base, while AS will probably wind up with its own special interest group (SIG).

References

[1]
Minutes of the AGATA Local Level Processing meeting, Liverpool, 24-25 February 2003.

[2]
IEEE 802.3z / IEEE 802.3ab specifications

[3]
M. Bellato et al “The CMS Event Builder Demonstrator based on GigaEthernet Switched Network”, CHEP 2000 Conference, http://chep2000.pd.infn.it/pres/pre_b307.pdf
[4]
CMS TRIDAS Project, Technical Design Report, vol. 2: “Data Acquisition and High Level Trigger”,


December 15 2002, CERN/LHCC 02-26 CMS TDR 6, also available  at http://cmsdoc.cern.ch/cms/TDR/DAQ/daq.html
[5]
PCI-Express specifications v. 1.0

� This was changed in draft 3 to allow for full rate data collection in some cases.


� Assume that the gammas detected in the core will each undergo 2 Compton-scatters and a photoevent (or 3 Compton scatters), so total count rate is 50kHz x 3 active segments to be processed. In addition, traces from 4 neighbours of each active segment will be sent to PSA. So, assuming no segment overlap, the maximum output data rate is: (1+ (3 x 5)) x 50kHz x 200bytes = 160Mbytes/second per crystal.





PAGE  
1

[image: image26.wmf]Pre

-

process:

E, BLR (+?)

Concentrator

(37 in, 1 out) 

with 

zero suppression

FADC data

from fibres

(36 segments)

Pre

-

process:

E, BLR 

+ local trigger

Global

Trigger I/f

FADC data

from fibre

(1 = core)

FADC Clock &

Commands to

digitiser 

Fibre or 

PCI

-

express link 

To PSA

To/from

GTS

[image: image27.wmf]AGATA Pre-processing costs. All prices Exclude taxes (VAT)

Contingency

20%

Initial Prototype

ATCA Prototype (1 Crystal)

Qty

unit price

cost

Segment Mezzanine (1st)

1

€ 2,813

€ 2,813

Segment Mezzanine (remainder)

5

€ 2,381

€ 11,904

Core mezzanine

1

€ 2,782

€ 2,782

GTS Mezzanine

1

€ 2,640

€ 2,640

CPCI Carrier

2

€ 4,470

€ 8,940

ATCA crate, switches, CPU

1

€ 20,000

€ 20,000

Cost without NRE

€ 49,078

Price per channel

€ 1,326

NRE (tooling)

Segment Mezzanine

1

€ 1,950

€ 1,950

Core Mezzanine

1

€ 1,950

€ 1,950

GTS Mezzanine

1

€ 6,000

€ 6,000

ATCA Carrier

1

€ 3,720

€ 3,720

Contingency on NRE

1

€ 2,724

€ 2,724

Total cost of NRE (tooling)

€ 16,344

Total cost of prototype (1 crystal)

€ 65,422

Development Costs

Cores, IP, development kits etc.

€ 100,000

(one off development cost)

Demonstrator

Production system (ATCA)

Price for 1 Cluster (3 crystals)

Segment Mezzanine 

18

€ 2,262

€ 40,712

Core mezzanine

3

€ 2,643

€ 7,928

GTS Mezzanine

3

€ 2,508

€ 7,524

ATCA Carrier

6

€ 3,766

€ 22,594

ATCA crate and switches

0.5

€ 20,000

€ 10,000

Cost without NRE (ATCA, 3 crystals)

€ 88,757

Cost per crystal

€ 29,586

price per channel

€ 800

NRE (tooling)

Rework of ATCA Carrier (50% of carrier NRE)

1

€ 1,860

€ 1,860

Rework of mezzanines (50% mezz NRE)

1

€ 4,950

€ 4,950

Contingency on NRE

1

€ 1,362

€ 1,362

Total cost of NRE (tooling)

€ 8,172

Total cost of first ATCA cluster incl NRE

€ 96,929

No spares included- assume the prototype is used for spares.

